Data Flow System

Document Title:	VISTA Infra Red Camera Calibration Plan

Document Number:			VIS-SPE-IOA-20000-0002

Issue:					0.5

Date:					2004-04-08

| Document Prepared by: | Peter Bunclark (CASU) | Signature And date: |  |
| --- | --- | --- | --- |
| Document Approved by: | Mike Irwin (CASU Manager) | Signature And date: |  |
| Document Released by: | Jim Emerson (VDFS Project leader) | Signature And date: |  |
| Document Reviewed by: | William Sutherland (VISTA Project Scientist) | Signature And date: |  |

# Change Record

| Issue | Date | Sections affected | Remarks |
| --- | --- | --- | --- |
| 0.5 | 2004-04-08 | All | New Document |

# Notification List

The following people should be notified by email that a new issue of this document is available.

| IoA: | W Sutherland |
| --- | --- |
| Oxford: | G Dalton |
| QMUL | J Emerson |
| ATC | Malcolm Stewart Steven Beard |

Table of Contents

Change Record	2

Notification List	2

1	Introduction	5

1.1	Purpose	5

1.2	Scope	5

1.3	Applicable Documents	5

1.4	Reference Documents	6

1.5	Abbreviations and Acronyms	6

1.6	Glossary	7

2	Overview	8

3	Observing Modes	12

3.1	IMAGING Mode	12

3.1.1	Description	12

3.1.2	Calibrations	12

3.2	High Order Wave Front Sensor (HOWFS) Mode	12

3.2.1	Description	12

3.2.2	Calibrations	13

4	Calibration Data for Instrumental Signature Removal	13

4.1	Purpose	13

4.2	Reset Frames	14

4.3	Dark Frames	14

4.4	Dome flats	15

4.5	Detector Noise	16

4.6	Linearization Measurements	16

4.7	Twilight Flats	17

4.8	Illumination Correction Measurement	18

4.9	Image Persistence Measurements	19

4.10	Electrical Cross-Talk Measurements	19

5	Data for Photometric Calibration	20

5.1	Introduction	20

5.2	Photometric Standards	21

5.3	Apply Photometric Calibration	22

6	Calibration Data Derived from Science Data	23

6.1	For Instrument Signature Removal	23

6.1.1	Night-Sky Flats	23

6.1.2	Sky Subtraction and Fringe Removal	23

6.1.3	Jittering	24

6.1.4	Microstepping	25

6.2	For Astrometric Calibration	25

6.2.1	Optical Distortion Effects	25

6.2.2	Final WCS Fit	26

7	Quality Control	27

7.1	Further Quality Control Data Derived from Science Frames	27

7.1.1	Object Extraction	27

7.2	On line quality control (QC-0)	27

7.3	Trend Analysis (QC-1)	28

8	Templates	30

8.1	Calibration Templates	30

8.1.1	IMAGING mode exposures	30

8.1.2	HOWFS mode exposures	32

8.2	Science Templates	32

8.2.1	IMAGING mode exposures	32

8.2.2	HOWFS mode exposures	40

8.3	Instrument Health Templates	41

9	Technical Programs	41

9.1	TP-VIS1: Establishment of Secondary Standard Fields	41

10	Format of Data Frames	43

10.1	Principle	43

10.2	Model FITS header	43

Appendix A.	All-Sky IR Survey Fields	48

# Introduction

## Purpose

This document forms part of the package of documents for the Preliminary Design Review of the Data Flow System for VISTA, the Visible and Infrared Telescope for Astronomy. As stated in [AD1] “The Calibration Plan is the prime document which describes the different instrument-specific components of the Data Flow System” and at PDR consists of a “draft version of the Calibration Plan including a list of all planned templates.” It will be developed from an incomplete form at PDR to a complete form at the FDR.

## Scope

This document describes the VISTA DFS calibration plan for the output from the 16 Raytheon VIRGO IR detectors in the (Infra Red) camera for VISTA.  The baseline requirements for calibration are included in the VISTA DFS User Requirements [AD2]. The major reduction recipes and algorithms to be applied to the data are described in the VISTA DFS Data Reduction Specification [RD1].

Each camera exposure will produce a ‘pawprint’ consisting of 16 non-contiguous images of the sky, one from each detector. The VISTA DFS will remove instrumental artefacts, combine the pawprint component exposures offset by small jitters, and photometrically and astrometrically calibrate each pawprint. It will also provide Quality Control measures. It will not combine multiple adjacent pawprints into contiguous filled images, nor stack multiple pawprints at the same sky position.

This document does not describe any calibrations or procedures relating to the CCD detectors that are also located within the camera and which interact with the Telescope Control System.

This document covers only the Routine Phase of operations of VISTA’s IR Camera. In particular it does not describe any calibrations or procedures that form part of the Commissioning Plan for VISTA, nor any procedures needed during routine Engineering Maintenance. [An exception is VIRCAM_howfs_obs_exp which is included here as it uses the science detectors.] Arrangements for processing any calibrations or procedures carried out under such categories are the responsibility of the VISTA Project Office.

## Applicable Documents

The following documents, of the exact issue shown, form part of this document to the extent specified herein. In the event of conflict between the documents referenced herein and the contents of this document, the contents of this document shall be considered as a superseding requirement.

- Data Flow for the VLT instruments requirements specification, VLT-SPE-ESO-19000-1618, issue 1.0, 1999-04-21.

- VISTA Infra Red Camera DFS User Requirements, VIS-SPE-IOA-20000-00001, issue 0.5, 2004-04-08.

## Reference Documents

The following documents are referenced in this document.

- VISTA Infra Red Camera Data Reduction Specifications, VIS-SPE-IOA-20000-00003, issue 0.5, 2004-04-08.

- Data Interface Control Document, GEN-SPE-ESO-19940-794, issue 2.0, 2002-05-21.

- VISTA Operational Concept Definition Document, VIS-SPE-VSC-00000-0002 issue 1.0, 2001-03-28

- VISTA Infrared Camera Technical Specification, VIS-SPE-ATC-06000-0004, issue 2.0, 2003-11-20

- VISTA IR Camera Software Functional Specification, VIS-DES-ATC-06081-00001, issue 2.0, 2003-11-12.

- IR Camera Observation Software Design Description, VIS-DES-ATC-06084-0001, issue 2.0, 2004-04.

- VISTA Science Requirements Document, VIS-SPE-VSC-00000-0001, issue 2.0, 2000-10-26

- A New System of Faint Near-Infrared Standard Stars, Persson et al., Astrophys. J. 116, 2475-2488, 1998

- JHΚ standard stars for large telescopes: the UΚIRT Fundamental and Extended lists, Hawarden et al., Mon.Not.R.Soc. 325, 563-574,2001

- The FITS image extension, Ponz et al, Astron. Astrophys. Suppl. Ser. 105, 53-55, 1994

- Representations of world coordinates in FITS, Griesen, & Calabretta,  A&A, 395, 1061.2002

- Representations of celestial coordinates in FITS, Calabretta & Griesen,  A&A, 395, 1077, 2002

- Overview of VISTA IR Camera Data Interface Dictionaries, VIS-SPE-IOA-20000-0004, 0.1, 2003-11-13

- Northern JHΚ Standard Stars fro Array Detectors, Hunt et al Astr.J 115, 2594, 1998

## Abbreviations and Acronyms

CDS			Correlated Double Sampling

DAS			Data Acquisition System

DFS 			Data Flow System

IMPEX		Import Export (P2PP ASCII files)

HOWFS		High Order Wave-Front Sensor

ICRF			International Coordinate Reference Frame

IWS			Instrument Workstation

FITS			Flexible Image Transport System

LOWFS		Low Order Wave-Front Sensor

OB			Observation Block

OT			Observing Tool

QC-0			Quality Control, level zero

QC-1			Quality Control, level one

TCS			Telescope Control System

URD			User Requirements Document

VDFS 			VISTA Data Flow System

VIRCAM		VISTA Infra Red Camera

VISTA			Visible and Infrared Survey Telescope for Astronomy

VPO			VISTA Project Office

WFCAM		Wide Field Camera (on UΚIRT)

WCS			World Coordinate System

2MASS		2 Micron All Sky Survey

## Glossary

To aid the understanding of the concepts in logical order the glossary is not alphabetical.

Confidence Map	An integer array, normalized to a median of 100% which is associated with an image. Combined with an estimate of the sky background variance of the image it assigns a relative weight to each pixel in the image and automatically factors in an exposure map.  Bad pixels are assigned a value of 0. It is especially important in image filtering, mosaicing and stacking.

Exposure	The stored product of many individual integrations, that have been co-added in the DAS. Each exposure is associated with an exposure time.

Integration	A simple snapshot, within the DAS, of a specified elapsed time. This elapsed time is known as the integration time.

Jitter (pattern)	A pattern of exposures at positions each shifted by a small movement (<30 arcsec) from the reference position. Unlike a microstep the non-integral part of the shifts is any fractional number of pixels. Each position of a jitter pattern can contain a microstep pattern.

Preset	A telescope slew to a new position involving a reconfiguration of the telescope control system and extra housekeeping operations that are not necessary for a movement or an offset.

Microstep (pattern)	A pattern of exposures at positions each shifted by a very small movement (<3 arcsec) from the reference position. Unlike a jitter the non-integral part of the shifts are specified as 0.5 of a pixel, which allows the pixels in the series to be interleaved in an effort to increase resolution. A microstep pattern can be contained within each position of a jitter pattern.

Mesostep	A sequence of exposures designed to completely sample across the face of the detectors in medium-sized steps to monitor residual systematics in the photometry.

Movement	A change of position of the telescope that is not large enough to require a new guide star.

Offset	A change of position of the telescope that is not large enough to require a telescope preset, but is large enough to require a new guide star.

Pawprint	16 non-contiguous images of the sky produced by the VISTA IR camera, with its 16 non-contiguous chips (see Figure 2-2). The name is from the similarity to the prints made by the padded paw of an animal (the analogy suits earlier 4-chip cameras better).

Tile	A filled area of sky fully sampled (filling in the gaps in a pawprint) by combining multiple pawprints. Because of the detector spacing the minimum number of pointed observations (with fixed offsets) required for reasonably uniform coverage is 6, which would expose each piece of sky, away from the edges of the tile, to at least 2 camera pixels. The pipeline does not combine pawprints into tiles.

# Overview

VISTA is a wide field alt-az telescope designed for a single purpose, surveys, and which does not have a conventional focus. It can only be used with a purpose built camera, and is delivered with an IR camera. Thus it is the performance and pointing of the telescope-camera system that is important.

The telescope by itself has no capability to lock onto a guide star or carry out wave front sensing. The IR Camera therefore contains, as well as 16 IR detectors, two Autoguider CCDs, and two low order wave front sensor (LOWFS) units, each with two CCDs, operating in the I band, as shown in Fig 2-1. Two autoguiders, on opposite edges of the focal plane, are used in order to meet the sky coverage requirements, although only one is allowed to apply corrections to the telescope axes at any given time. The LOWFSs measure aberrations that are used by the external active optics control process to adjust the position of the 5 axis (x, y, z, tip, tilt) secondary mirror support system and some aspects of the M1 surface to maintain image quality. The LOWFS operates roughly every 1 minute during tracking and needs exposures of ~40 sec to average out seeing effects. Although the Autoguiders and LOWFSs are physically located within the IR camera, both are considered part of the TCS from a software point of view. This is primarily to maintain consistency with existing VLT software and standards. The VISTA DFS receives no data from these CCDs. The CCDs therefore do not impact on the VISTA DFS, except in so far as the pointing and image quality of the camera are dependent on their proper operation.

A high order wave front (curvature) sensor (HOWFS) uses some of the science detectors to determine occasional adjustments to the primary mirror support system. (This is done perhaps once at the start of the night and once around midnight.) Processing the signals from the HOWFS is done within the Instrument Workstation, and so the pipeline will not have to deal with the HOWFS at all. However all data from the IR detectors, including HOWFS data, is passed to the science archive, so the necessary calibration templates for the HOWFS are covered here.

+Y

+X

Figure 2-1 VISTA Focal plane: Each of the 4 groups of detectors in the Y direction (e.g. #s 1-4, 5-8, 9-12, 13-16) is read out by a separate IRACE controller.

Within the IR Camera are 16 Raytheon 2048x2048 VIRGO detectors arranged in a sparse array. Each camera exposure produces a pawprint consisting of 16 non-contiguous images of the sky. An example display of a complete FITS file consisting of a (synthesized) VISTA “pawprint” is shown in Figure 2-2.

The VISTA IR camera has only one moving part, the filter wheel which has 8 filter holders, each filter holder containing 16 filters, one for each IR detector. There are further auxiliary (beam splitting) filters for use with the high order wave front sensor. One of the filter holders contains a set of 16 cold blanks (metal units which completely block the detectors from incoming sky radiation, and produce negligible thermal emission) which are used for taking dark frames. The instrument will be delivered with 4 filter sets (Y, J, H, Κs) and a further three sets of cold blanks, which can be replaced with other filters in due course. The position angle of the camera axis can be controlled by the instrument rotator. Single integrations are taken by a Reset-Read-Read procedure with the difference of the two Reads being performed within the DAS.

The VISTA DFS will produce photometrically and astrometrically calibrated pawprints, with instrumental artefacts removed. In order to achieve almost uniform coverage of a full contiguous area of sky, a six point offset pattern is used by default. A template that implements this pattern is defined and the DFS will calibrate the resulting six pawprints individually.  The further step of combining these into a contiguous map is left to the science user.

Other processes which are not calibration issues, but which may nevertheless relate to achievable data quality, are not discussed here. Such (excluded) processes include:

- co-addition of individual integrations of a pawprint into a single exposure within the data-acquisition system;

- combination of many pawprints to cover contiguous areas of sky;

- co-addition of many pawprints to go deeper.

Section 3 defines the observing modes, Section 4 describes the calibration data required for instrumental signature removal, Section 5 describes the calibration data required for photometric calibration. Section 6 describes the calibration data to be derived from science data, including astrometric calibration. Section 7 discusses Quality Control measures based on regularly measured selected sets of calibrations for the purpose of instrument “health checks”. Section 8 describes all templates and Section 9 the Technical Programs. Finally Section 10 details the Format of Data Frames.

The philosophy throughout is that the DFS pipeline will be triggered by the ‘arrival’ of each exposure, and by the completion of a template (when one might want to operate on many of the exposures gathered during the template). The content of the FITS headers, for example OB identifiers, can allow the DFS pipeline to handle the set of observed files from the OB as an ensemble, and to choose appropriate processing based on the header information.

Figure 2-2 Synthesized VISTA Pawprint.

# Observing Modes

## IMAGING Mode

IMAGING is the only mode in which science data will be acquired.

### Description

The sky target position is acquired and tracked and in parallel (for observing efficiency) the required filter set is placed in the beam. The LOWFS provides the necessary updates to the M2 and M1 support units. A set of exposures, each of which may consist of a number of integrations, are taken and are usually jittered by small offsets, to remove bad pixels and determine sky background. The set of exposures produced is combined in the DFS to create a single pawprint, in which the jitters from all detectors are included.

Six such pawprints, taken at appropriate offsets, can be combined to produce an almost uniformly sampled image of a contiguous region, each bit of sky, except at the edges, having been observed by at least two pixels. The individual exposures making up each pawprint may be made on a jitter or a microstep pattern. Microstep patterns are interleaved rather than combined, so the calibration procedures are unchanged, though the data volume increases.

### Calibrations

The calibrations are of four sorts:

- those that characterize the properties of the transfer function (image in, electrons out) of the end-to-end system (telescope, camera, IR detector system including associated controllers, etc.) so that instrumental effects can be removed from the data. As VISTA has a wide field of view, particular attention must be paid to variations across the field;

- those that characterize the astrometric distortions of the images;

- those that characterize the photometric zero points and extinction coefficients corresponding to the images;

- those that generate Data Quality measures.

## High Order Wave Front Sensor (HOWFS) Mode

The HOWFS mode is processed in the Instrument Workstation and is logically part of the TCS. However, as it uses the IR detectors, all of whose data is passed to the archive, it is considered as a separate observing mode for VISTA DFS purposes.

### Description

In HOWFS mode a special beam splitting filter is used to make a curvature sensor in which two images (above and below focus) of a reference star are formed and used to generate corrections to the forces in the M1 support unit, ensuring the mirror figure is maintained. This mode will typically be used of order twice a night (start and around midnight), or less often if the repeatability of the lookup table is good.

### Calibrations

The HOWFS uses some of the science mode IR detectors, but has a special beam splitting filter. Thus its flat-field needs to be calibrated separately from the science mode IR detectors.

# Calibration Data for Instrumental Signature Removal

## Purpose

Section 4 describes what calibration data has to be collected with what frequency to allow one to remove instrumental signatures.

For each piece of calibration data required this section defines:

- Responsible: responsibility for obtaining the calibration data

- Phase: when the calibration data has to be acquired (day or night time)

- Frequency: how often calibration data need to be acquired.

- Purpose: reason for needing the calibration data

- Procedure: the procedure for acquiring the calibration data

- Raw Outputs: the output of the procedure (number of FITS files by FDR)

- Prepared OBs/Templates: the pre-prepared observation blocks or templates to acquire the calibration data

- OT queue: the corresponding Observing Tool queue for the Observation Blocks.

- Pipeline Recipe: The name (if any) of the processing recipe applied by the data flow system pipeline. Recipes may contain algorithms and procedures as subcomponents. Each such recipe corresponds to one listed in [RD1]. The exact specification of each recipe, and its subcomponents, will appear in later issues of [RD1] as only the list itself is required for PDR

- Pipeline Output: the Pipeline output products, appended with (QC) for those also used as Quality Control parameters

- Duration: an estimate of the required time to execute the calibration procedure (time to process the raw output in DFS pipeline will be added by FDR)

- Prerequisites:  possible dependencies on instrumental or sky conditions or other calibration procedures are given

- See also: any further information.

The calibration data is used for instrumental signature removal. The aim is to provide pawprints as though taken with a perfect camera, which produces a photometrically linear, defect-free, evenly-illuminated reproduction, though sparsely sampled, reproduction of the sky. This will have no additional systematic, random noise or other artefacts, and will be on an arbitrary photometric and astrometric scale.

Off-sky calibrations and quality control measures will be made routinely, before and after observing, using the in-dome illuminated screen.

## Reset Frames

| Responsible: | Science Operations |
| --- | --- |
| Phase: | Daytime |
| Frequency: | Daily |
| Purpose: | A Reset frame is a Reset-Read sequence with minimum exposure (1 sec is the minimum VISTA can produce) taken with the cold blank in. It differs from a dark frame, which consists of a Reset-Read-Read sequence where the output is the difference of the two reads. The aim is to map the effect of the reset. Sequences of Reset frames will be taken off-sky and analysed to estimate the stability of the reset pedestal and pixel to pixel variation. |
| Procedure: | Read out frame, compare with library reset frame. |
| Raw Outputs: | FITS files |
| Template: | VIRCAM_img_cal_reset.tsf |
| OT queue: | VIRCAM.Daytime.Calibration |
| Pipeline Recipe: | vircam_reset_combine |
| Pipeline Outputs: | Variance with respect to standard frame (QC) |
| Duration: | 1 s |
| Prerequisites: |  |
| See Also: |  |

## Dark Frames

| Responsible: | Science Operations |
| --- | --- |
| Phase: | Daytime |
| Frequency: | Daily |
| Purpose: | Dark Frames are used to calibrate out two separate additive effects. - the accumulated counts that results from thermal noise (dark counts). This is generally a small, but not negligible effect. - an effect, here called ‘reset anomaly’, in which a significant residual structure is left in the image after the reset is removed in the DAS, when it does a correlated double sample (CDS, Reset-Read-Read). Both dark counts and reset anomaly are additive and can be removed together, using dark frames (exposures with cold blank filters completely blocking the detectors from incoming radiation) taken with the same integration time as the target observation. In order to minimize contamination from transient events, a dark frame would be a combination of many frames with rejection. If the spatial structure of the reset anomaly is not stable with time it could leave a challenging background variation over the detector, which may need to be removed with a background filter. This latter scenario is best avoided as real astronomical signal will inevitably be removed. (In general, for other instruments examined where the reset anomaly structure is repeatable and stable, the integration time seems to determine the spatial structure of the residuals, while the ambient flux seems to determine its intensity.) |
| Procedure: | A series of dark frames will be taken with each integration and exposure time combination used for target observations so that the structure of the reset anomaly can be modelled correctly and the dark count correction is consistent. The Dark template will not move the telescope, will insert the cold blank and perform a timed exposure. If the requested time is less than the array minimum read-out cycle time of ~1s (e.g. zero) the controller will deliver, and report, the minimum detector integration time of ~1s. |
| Raw Outputs: | FITS Files |
| Template: | VIRCAM_img_cal_dark.tsf |
| OT queue: | VIRCAM.Daytime.Calibration |
| Pipeline Recipe: | vircam_dark_combine |
| Duration: | One for each integration and exposure setting for the science observations made on the same night |
| Pipeline Outputs: | Master Dark Dark + reset anomaly stability trend (QC) Detector dark current (QC) Detector Particle Event rate (QC) |
| Prerequisites: |  |
| See Also: |  |

## Dome flats

| Responsible: | Science Operations |
| --- | --- |
| Phase: | Daytime or non-observing nights. |
| Frequency: | Daily |
| Purpose: | Monitoring instrument performance, image structure, and high-accuracy bad pixel mapping. They will not be used for gain correction (flat-fielding) due to non-uniform illumination over the whole of the focal plane and the different colour of the illumination compared to the night sky. |
| Procedure: | The Dome template will acquire the dome screen (constant illumination); a series of timed exposures are made through a given filter. |
| Raw Outputs: | FITS files |
| Prepared OBs: | VIRCAM_img_cal_domeflat.obx |
| OT queue: | VIRCAM.Daytime.Calibration |
| Pipeline Recipe: | vircam_dome_flat_combine |
| Pipeline Outputs | Updated Master dome flats Updated Bad pixel maps Bad pixel statistics (QC) |
| Duration: | 10 min |
| Prerequisites: | The need for constant illumination of the dome screen implies that the dome flats cannot be taken in conditions of variable ambient light; nor can they be taken in conditions of excessive ambient light. |
| See Also: | Dome flat observations are also employed in linearization measurements described in 4.6. |

## Detector Noise

| Responsible: | Science Operations |
| --- | --- |
| Phase: | Daytime |
| Frequency: | Daily |
| Purpose: | In order to understand the noise properties of the detectors, it is important to measure the readout noise and gain of each chip. This is a vital piece of information, not only as large changes in either property could signal a detector health issue, but also as further down the pipeline the issue of pixel rejection algorithms becomes important (for example, during jittering). |
| Procedure: | Both of these properties can be measured from a pair of dark exposure frames and a pair of dome flat frames. The dark exposures should have matching integration and exposure times to the dome flats, and should be observed with the same dome illumination. Care should be taken to ensure that the flats are exposed in a region of the response curve where the detectors are reasonably linear. |
| Raw Outputs: | FITS files |
| Prepared OBs: | VIRCAM_img_det_noise.obx |
| OT queue: | VIRCAM.Daytime.Calibration |
| Pipeline Recipe: | vircam_detector_noise |
| Pipeline Outputs: | Readout noise and gain estimate for each sector of each detector (QC) |
| Duration: | 1 minute |
| Prerequisites: |  |
| See Also: |  |

## Linearization Measurements

| Responsible: | Science Operations |
| --- | --- |
| Phase: | Daytime or cloudy nights (better) |
| Frequency: | Monthly |
| Purpose: | Infrared detectors can be strongly non-linear. The linearity curve of each detector can be determined through a series of differently timed dome screen observations under constant illumination. These curves are used to generate a look-up table of observed-to-true data values which can be applied to the images in the pipeline. [Linearization within the DAS, which is an alternative, is not included in its Technical Specification]. |
| Procedure: | On a series of specified dates (monthly) take series of dome flats under constant illumination at varying exposures up to full counts. |
| Raw Outputs: | FITS files |
| Prepared OBs: | VIRCAM_img_cal_linearity.obx |
| OT queue: | VIRCAM.Daytime.Calibration |
| Pipeline Recipe: | vircam_linearity_analyse |
| Pipeline Output: | Linearization curve and lookup tables Trend of non-linearity function(QC) |
| Duration: | [30] min |
| Prerequisites: | The need for constant illumination of the dome screen implies that the dome flats cannot be taken in conditions of variable ambient light; nor can they be taken in conditions of excessive ambient light. |
| See Also: | Dome flat measures in 4.4 |

## Twilight Flats

| Responsible: | Science Operations |
| --- | --- |
| Phase: | Twilight |
| Frequency: | Evening/Morning |
| Purpose: | Flat-fielding removes multiplicative instrumental signatures from the data. This includes pixel-to-pixel gain variations and the instrumental vignetting profile. It also provides a global gain correction between detectors and individual read out channels within each detector. (Each of the 16 detectors has 16 read out channels, giving a total of 256 read out channels.) Mean flat-fields also are the data source for the initial confidence map for each detector and filter combination. This is similar to a weight/bad-pixel map where the mean level is normalized to a value of 100% and bad pixels are flagged with a value of zero. A new confidence map is formed each time some sort of image combination is done. It is used in conjunction with an estimate of the sky background variance in each frame to propagate the weight of each individual pixel. Although this is especially important for later manipulation of the pawprints outside the VISTA DFS for doing deep stacking and tiling, it is also vital for the object detection part of the pipeline which is used, inter alia, in astrometric and photometric corrections. Mean flat-fields can be derived from a variety of sources (each with their own advantages and disadvantages). Sky flats taken at twilight have a good (but not perfect) colour match to the night sky observations we wish to correct, and can be taken under conditions where the contribution from night sky fringing, emission from dust (on the optical surfaces) and other spatial effects are most negligible. There will still be a very small residual error in gain correction as the twilight sky has a slightly different colour to the night time sky. Dusk and dawn twilight flats will be combined to moderate effects caused by the significant variation in the illumination caused by the reset and read times. |
| Procedure: | The sky level must be such that any emission from fringing or dust on the optical surface will be negligible in comparison, and this means that there is only a short time in which to acquire the twilight flats. It will not always be possible to get a complete set of twilight flats every night for schedules involving many filters or on nights with changeable weather. If, however, the detector flat-fields are sufficiently stable, then it is possible to use master flats taken over several nights, which is the method of choice. |
| Raw Outputs: | FITS Files |
| Prepared OBs: | VIRCAM_img_cal_twiflat.obx |
| OT queue: | VIRCAM.Daytime.Calibration |
| Pipeline Recipe: | vircam_twilight_combine |
| Pipeline Output: | Master twilight flats Initial confidence maps Trend in mean gain correction between detectors (QC) |
| Duration: | 10 min evening twilight, 10 min morning twilight. |
| Prerequisites: |  |
| See Also: |  |

## Illumination Correction Measurement

| Responsible: | Science Operations |
| --- | --- |
| Phase: | Night |
| Frequency: | Monthly |
| Purpose: | The gain correction as modelled by the flat-field should remove all pixel-to-pixel gain differences as well as any large-scale variations due (generally) to vignetting within the focal plane. However, scattered light within the camera may lead to large-scale background variations which cannot be modelled and removed, as its level depends critically on the ambient flux. Dividing a target frame by a flat-field frame that is affected by this will cause systematic errors in the photometry across the detector. It is necessary to map out the spatial systematic effects across each detector so that a correction map can be factored into the final photometry measured from each detector. |
| Procedure: | The illumination correction can be measured in two ways. In the event that observations of a secondary photometric standard field with a density of 100-200 objects per detector are available, then the illumination correction can be measured by looking at the spatial variation of the photometric zero-point across each detector. If such a field is not available, then a mesostep sequence is taken consisting of a series of exposures of a sparse field of relatively bright stars on a regular grid of offsets that cover one detector. Measuring a flux on each exposure allows the definition of a position-dependent scale factor (this must be done for each filter and each detector). |
| Raw Outputs: | FITS files |
| Prepared OBs: | VIRCAM_img_cal_illumination.obx |
| OT queue: | VIRCAM.Nighttime.Calibration |
| Pipeline Recipe: | vircam_illumination_analyse, vircam_mesostep_analyse |
| Pipeline Output: | Correction map |
| Duration: | 30 min |
| Prerequisites: |  |
| See Also: |  |

## Image Persistence Measurements

| Responsible: | Science operations |
| --- | --- |
| Phase: | Night |
| Frequency: | Monthly, and on detector/controller change |
| Purpose: | Image persistence (sometimes also called ‘remanence’) is the effect where residual impressions of images from a preceding exposure are visible on the current image. |
| Procedure: | On a sequence of (monthly) dates choose a fairly empty field with a nearly saturated star. Take an exposure and then a sequence of dark frames to measure the characteristic decay time. This must be done for each detector. |
| Raw Outputs: | FITS files |
| Prepared OBs: | VIRCAM_img_cal_persistence.obx |
| OT queue: | VIRCAM.Nighttime.Calibration |
| Pipeline Recipe: | vircam_persistence_analyse |
| Pipeline Output: | Decay constant Trend of decay constant and possible adjacency effects (QC) |
| Duration: | 10 min |
| Prerequisites: |  |
| See Also: |  |

## Electrical Cross-Talk Measurements

| Responsible: | Science operations |
| --- | --- |
| Phase: | Night |
| Frequency: | Monthly |
| Purpose: | Electrical cross-talk will be measured in the laboratory and during commissioning, and is expected to be negligible. As cross-Talk might change with any changes to the electrical environment a routine procedure to check it is planned. |
| Procedure: | The 16 detectors are read out in 16 sectors, making a total of 256 sectors in the camera. Cross-Talk calibration consists of placing a saturated star on a sector and measuring any effect on the other 255 sectors. This results in a 256x256 matrix, the majority of whose elements will hopefully be zero. Any electrical cross Talk between different detectors is anticipated to be smaller than between sectors within a detector. No specific template is necessary, as a suitably crafted observation block will be used. |
| Raw Outputs: | FITS Files |
| Prepared OBs: | VIRCAM_gen_tec_crosstalk.obx |
| OT queue: | VIRCAM.Nighttime.Calibration |
| Pipeline Recipe: | vircam_crosstalk_analyse |
| Pipeline Output: | Cross-talk matrix. Trend of off-diagonal components (QC) |
| Duration: | 10 min |
| Prerequisites: |  |
| See Also: |  |

# Data for Photometric Calibration

## Introduction

The camera will be on the telescope semi-permanently, in a survey mode, providing a stable configuration which enables a long-term approach to photometric calibration to be taken.  The strategy is to define routine calibration procedures, so that the accuracy, and hence the scientific value, of the archive, will be maximized. Magnitudes will be calibrated on the Vega scale.

At any time (t) on any night (n) for any star (i) in any filter waveband (b),

mcalib= minstibtn+ZPbtn−κbtn(X−1)				Equation 1

where ZP is the Zero Point (i.e. the magnitude at airmass unity which gives 1 count/second at the detector), mcal is the calibrated instrumental magnitude, minst is the measured instrumental magnitude (-2.5 × log10[counts/sec]), κ is the extinction coefficient and X is the airmass of the observation.  This assumes that the second-order extinction term and colour-dependency of κ are both negligible.

Typically, the Zero Point of the instrument + telescope system should be stable throughout the night.  Long-term decreases in the sensitivity of the instrument, and hence a decreasing ZP, could be caused by for example the accumulation of dust on the primary mirror.

On photometric nights the extinction coefficient κ should be constant in each filter. The extinction κ will be monitored through each night assuming a fixed zero point and making measurements over a range of airmass. Although 2MASS found their extinction coefficients to vary seasonally any effect should be much less for VISTA since it has narrower filter profiles especially at J, and is at a much drier site.

A network of Secondary Standard photometric fields will be set up so that routine photometric standard observations can be made with the telescope in focus every hour. Frequent time-sampling is required to ensure accurate measurement of extinction and Zero Point, and so that the photometricity of a night can be monitored and derived from the data. Many of these standard fields (the equatorial ones) will have been observed and calibrated in advance by WFCAM.  The secondary fields will meet the following criteria:

- Extend over the area of the IR camera pawprint

- Span 24 hours in RA, with a target spacing of 2 hours.

- Enable observations over a range of airmass. Therefore, some must be chosen to pass close to the zenith of VISTA (for airmass unity), and such that fields will be available to the north and south of the zenith to optimize telescope azimuth slewing.  The remainder will be near-equatorial.

- Have a density of sources sufficient to characterize the systematic position-dependent photometric effects in VISTA, but be not too crowded.  The target is of order 100 stars per detector with magnitudes no fainter than J=18, Κs=16 to avoid prohibitively long exposures.

- They should encompass as broad a spread as possible in colour in order to derive colour terms robustly and facilitate transformations from and to other filter systems and e.g. the AB magnitude system.  i.e.

Mstd=mbcal+C(Mxstd−Mystd) 				Equation 2

where Mstd is the magnitude in a defined standard system, mbcal is the calibrated magnitude in the instrumental system, and C is the colour term for the appropriate standard colour index(Mxstd−Mystd).

- Each field should be centred on a primary standard: either a UΚIRT Faint Standard [RD9] and www.jach.hawaii.edu/JACpublic/UΚIRT/astronomy] or a LCO/Palomar NICMOS (Persson) standard [RD8], or a standard from [RD14], enabling direct calibration of the secondary standards. There are sufficient of these faint primary standards so that we can select primary standards which will not saturate the detector in a short (seconds) in-focus exposure.

- Technical Program TP-VIS1 describes the observations needed to set up the secondary standard fields.

## Photometric Standards

| Responsible: | Science Operations |
| --- | --- |
| Phase: | Night |
| Frequency: | Hourly |
| Purpose: | Determine ZP and κ to allow application of mcalib= minstibtn+ZPbtn−κbtn(X−1) to photometrically calibrate all objects seen. |
| Procedure: | Suitable fields from this network will be observed over a range of airmass each night to determine the Zero Points (ZP) and to monitor the extinction coefficients (κ) for all broad-band filters. |
| Outputs: | FITS files |
| Prepared OBs: | VIRCAM_standard |
| OT queue: | Science |
| Pipeline Recipe: | vircam_photcal_fit |
| Pipeline Output: | Zero Point (ZP) Zero Point trend (QC) Extinction coefficient (κ) Extinction coefficient trend (QC) Colour terms (C) Illumination correction Global gain correction (check) |
| Duration: | 5 min 10 times per night |
| Prerequisites: |  |
| See Also: |  |

## Apply Photometric Calibration

| Responsible: | Science Operations |
| --- | --- |
| Phase: | Night |
| Frequency: | All on sky data |
| Purpose: | Apply Photometric Calibration |
| Procedure: | Apply mcalib= minstibtn+ZPbtn−κbtn(X−1) using ZP and κ found from photometric calibration fields to calibrate frames photometrically. |
| Outputs: | Photometry FITS headers |
| Prepared OBs: | None |
| OT queue: |  |
| Pipeline Recipe: | vircam_photcal_apply |
| Pipeline Output: | Calibrated frames Depth of Exposure (QC) |
| Duration: |  |
| Prerequisites: |  |
| See Also: |  |

# Calibration Data Derived from Science Data

## For Instrument Signature Removal

### Night-Sky Flats

| Responsible: | Science Operations |
| --- | --- |
| Phase: | Night |
| Frequency: | Throughout night |
| Purpose: | In the situation where the detector flat-field is not stable over the course of a night, we will use night-sky flats. These are formed either from the target frames or from any special offset sky frames that might have been taken (for example where there is a large extended object in the field). All such frames over an appropriate time range are combined with rejection to form a normalized night sky flat-field. The advantage over twilight flats is the identical colour match between the sky observations and the targets, which means that there is no colour variation in the gain correction. However, fringing and thermal emission from dust particles on the optical surfaces can be high enough to affect the background significantly in some passbands. Dividing the target frames by a flat without correcting for these two additive effects could lead to significant systematic errors in photometry. In the Garching pipeline, master flats will be determined from as many observations as possible, but if it is determined that the flats vary rapidly, then only flats taken close in time may be useable. |
| Procedure: | Use normal science exposures. |
| Raw Outputs: | FITS Files |
| Prepared OBs: | None |
| OT queue: | science |
| Pipeline Recipe: | vircam_sky_flat_combine |
| Pipeline Output: | Night sky Flats Trend with respect to master flat (QC) Initial confidence map |
| Duration: | Occurs in parallel with all night observing |
| Prerequisites: | Determine the characteristics of fringing and thermal emission from dust on the optical surfaces during commissioning. |
| See Also: | 6.1.2 |

### Sky Subtraction and Fringe Removal

| Responsible: | Science operations |
| --- | --- |
| Phase: | Night |
| Frequency: | Throughout night |
| Purpose: | The sky background varies over large scales in the infrared. In some wavebands, fringing and thermal emission from any local dust (on optical surfaces) will also be present. All of these effects can be removed using the sky-subtraction algorithm. The source of the sky background estimate is usually the science data frames themselves. In cases where large extended or very bright objects might be present it may be necessary to use ‘offset sky’ exposures. |
| Procedure: | Preset or offset to, uncrowded, regions taken near or adjacent to the region of interest. Observe in the same way as the corresponding science field. |
| Raw Outputs: | FITS Files |
| Prepared OBs: | None |
| OT queue: | science |
| Pipeline Recipe: | vircam_offset_sky_combine |
| Pipeline Output: | Local sky estimate Fringe maps |
| Duration: | Same as science field. |
| Prerequisites: |  |
| See Also: |  |

### Jittering

| Responsible: | Science Operations |
| --- | --- |
| Phase: | Night |
| Frequency: | Nearly all the time |
| Purpose: | Removal of bad pixels and other cosmetic effects, as well as cosmic rays, and determining the sky background. Typically a long exposure is split into several shorter exposures, which, rather than being repeated with each pixel looking at exactly the same sky position, are carried out at a series of different (jittered) positions. This is similar to microstepping (same template), but with less fine sampling, and the pipeline combines the jittered exposures using a rejection algorithm. |
| Procedure: | Perform a specified pattern of exposures at each position of a jitter pattern. Predefined patterns and movement size in pixels may be selected. Microsteps can be nested within each jitter position by setting the number of microsteps appropriately in the template. |
| Raw Outputs: | FITS Files |
| Template: | VIRCAM_img_obs_pattern.tsf |
| OT queue: | Science |
| Pipeline Recipe: | vircam_jitter_combine |
| Pipeline Output: | Combined frames of pawprint Confidence map for pawprint |
| Duration: | Variable |
| Prerequisites: |  |
| See Also: | 6.1.4 |

### Microstepping

| Responsible: | Science operations |
| --- | --- |
| Phase: | Night |
| Frequency: | As required |
| Purpose: | Improved sampling. This is most likely to be employed in times of excellent seeing, when the point-spread function is under-sampled. It can also be used if there are strong intra-pixel sensitivity (QE) variations. It may not be commonly used. It is similar to jittering (same template) but with improved sampling through finer pattern spacing, and the pipeline interleaves the exposures without further rejection. |
| Procedure: | Perform a specified pattern of exposures at each position of a microstep pattern. Predefined patterns and movement size in pixels may be selected, and there is a default pattern/size [2×2 pattern, modulo a 0.5 pixel shift]. By setting the number of microstep appropriately in the template, microsteps can be nested within each jitter position. |
| Raw Outputs: | FITS Files |
| Template: | VIRCAM_img_obs_pattern.tsf |
| OT queue: | Science |
| Pipeline Recipe: | vircam_microstep_interleave |
| Pipeline Output: | Interleaved science frames with corresponding confidence maps |
| Duration: | Variable |
| Prerequisites: |  |
| See Also: | 6.1.3 |

## For Astrometric Calibration

Astrometric calibration will take the instrument signature free pawprints and provide the transformation between pixel coordinates and celestial coordinates for each of the 16 constituent images, though still leaving the pawprints on an arbitrary photometric scale. The transformations are manifested in a Flexible-Image Transport System (FITS) [RD10] World-Coordinate System (WCS) [RD12]. The projection used will be Zenithal Polynomial (ZPN), based on the predicted properties from the optical design.

Quantifying the distortion terms used in the WCS will be done from on-sky observations. An initial astrometric distortion is available from the optical design, and an updated early empirical value will be derived from commissioning data. Following that, an increasingly accurate value will be derived from the astrometry of all target frames.

### Optical Distortion Effects

| Responsible: | Science Operations |
| --- | --- |
| Phase: | Night |
| Frequency: | All science frames |
| Purpose: | The strongest term in the optical-distortion model is the cubic radial term, but this and all distortions will be slightly colour (i.e. filter) dependent and must be determined on sky. The expected power of the distortion means that no practically useful jitter is possible without non-linear resampling. The radial scale distortion also has an impact on photometric measurements, inducing an error up to 3.5% in the corners of the field, compared to the centre, if uncorrected. It is thus crucial to determine it accurately. |
| Procedure: | Astrometric stars in the science fields are used to map the distortion, an increasingly accurate description of which builds up from the astrometry of all target frames. |
| Raw Outputs: | FITS files |
| Prepared OBs: | None |
| OT queue: | Science |
| Pipeline Recipe: | vircam_distortion_update |
| Pipeline Output: | Refined optical distortion model |
| Duration: | No overhead |
| Prerequisites: | Initial value from optical design, an early empirical value from commissioning data, |
| See Also: |  |

### Final WCS Fit

| Responsible: | DFS calibration pipeline |
| --- | --- |
| Phase: | Night |
| Frequency: | All imaging frames on sky |
| Purpose: | The camera software writes an initial WCS based on the given position of the guide star (with accuracy better than 2˝, dependent on the guide star accuracy, and the determined geometry of the camera) into the FITS headers of each data frame. This provides a close starting point for orientation of the data frames and location of astrometric stars for a full WCS solution that will provide refined scientific quality astrometry. After instrumental-signature removal astrometric stars are centroided in the data frames to typically 0.1 pixels accuracy. An astrometric solution is carried out using reference catalogues based on the International Coordinate Reference Frame (ICRF) [e.g. 2MASS catalogue]. Accuracy is dependent on the reference catalogue accuracy, but the final uncertainty estimate comes from the RMS of the fit and the known systematics of the reference catalogue. |
| Procedure: | None |
| Raw Outputs: | None |
| Prepared OBs: | None |
| OT queue: | - |
| Pipeline Recipe: | vircam_wcs_fit |
| Pipeline Output: | Refined WCS FITS header for all frames Pointing accuracy (QC) [Calculated from equatorial coordinates computed at particular location using the fitted WCS and the initial WCS that was written to the raw header] |
| Duration: | Zero overhead |
| Prerequisites: | Commissioning to determine initial WCS |
| See Also: |  |

# Quality Control

## Further Quality Control Data Derived from Science Frames

### Object Extraction

| Responsible: | Science Operations |
| --- | --- |
| Phase: | Night |
| Frequency: | Nearly all the time |
| Purpose: | Object extraction is vital for various steps in the pipeline, including astrometric and photometric calibration, where the position and/or photometric measures of real objects are required. It is also needed in order to assess the quality of the data in terms of the observing conditions and the depth of exposure. |
| Procedure: | Extract objects from each frame using the object extraction algorithm. Classify objects as stellar, non-stellar and noise using the classification scheme. Use the stellar objects to work out the average properties of the images on the frame. |
| Raw Outputs: | FITS Files |
| Template: | - |
| OT queue: | Science |
| Pipeline Recipe: | vircam_gen_catalogue |
| Pipeline Output: | Mean sky background (QC) Mean sky noise (QC) Number of noise objects (QC) Mean seeing (QC) Mean stellar ellipticity (QC) |
| Duration: | Variable |
| Prerequisites: |  |
| See Also: |  |

## On line quality control (QC-0)

QC-0 is generic for all VLT-compliant instruments and is provided by the VISTA control software for VIRCAM. Therefore, the procedures described in this document are assured that they have been fed data which has passed QC-0.

## Trend Analysis (QC-1)

| Effect | Section | Description |
| --- | --- | --- |
| Mean reset frame variation | 4.2 | The current mean reset frame is compared with the most recent library frame. Subtracting one from the other should leave neither significant residuals nor any structure in the residuals map. |
| Reset anomaly variation | 4.3 | The current mean dark frame is compared with the most recent library frame that was taken with matching integration and exposure times. The difference of the two should yield neither significant residuals nor any structure in the residuals map. |
| Dark current trend | 4.2 | Comparing the current estimate with recent values for the same chip should show any systematic trends in dark current. |
| Particle Event Rate | 4.3 | Pixels rejected during the combination of dark frames will give an estimate of the rate of cosmic ray hits for each detector. This can be compared with previous estimates and monitored. |
| Bad pixel statistics | 4.4 | The fraction of bad pixels (either hot or cold) should not change significantly with time on a detector. |
| Detector readout noise and gain | 4.5 | The noise properties of each detector should remain stable so long as the electronics have not been modified. |
| Trend in non-linearity | 4.6 | Although all infrared detectors are non-linear to some degree, the shape and scale of the linearity curve for each detector should remain constant. New linearity curves should show no significant change from other recent estimates unless the electronics of the system have been modified. |
| Trend in mean gain correction | 4.7 | The ratio of median counts in a mean flat exposure for a given detector relative to that for another detector specifies the mean gain correction between those two detectors. The mean gain corrections for the ensemble should be very stable with time. |
| Trend in persistence decay time and adjacency effects | 4.9 | The image persistence decay constant for each detector should be monitored. |
| Trend in cross-talk matrix values | 4.10 | Values in the cross-talk matrix should be very stable with time, not withstanding an update of the driving electronics. |
| Variation in photometric zero point and extinction coefficient. | 5.2 | The photometric zeropoint will probably vary over time as a result of the cleaning of optical surfaces. Extinction will vary over a night, but wildly varying extinction coefficients are an indication of bad photometric conditions. |
| Depth of exposure | 5.2 | This can be compared with a target input value to see if the image has been exposed for long enough. |
| Variation of master flats | 6.1.1 | Dividing the current flat-field by a recent library flat-field for the same filter should leave neither significant residuals nor large scale variations. |
| Pointing accuracy | 6.2.2 | The equatorial coordinates of a location in the focal plane computed from the fitted WCS and computed from the raw file WCS should agree to < 2 arcseconds. |
| Trend in mean sky and sky noise | 7.1.1 | These will vary over the night, but should do so smoothly. If this is not the case, then we are observing in non-photometric conditions. A strange value here may also indicate a serious electronic fault. |
| Number of noise objects | 7.1.1 | The number of noise objects from frame to frame should be reasonably constant. |
| Mean seeing | 7.1.1 | The seeing will obviously vary over the night with time and wavelength (filter). This variation should be smooth. Wild variation could be an indication of non-photometric seeing. |
| Mean stellar ellipticity | 7.1.1 | Mean stellar ellipticity can be affected by failures in the tracking and autoguiding. |

# Templates

## Calibration Templates

### IMAGING mode exposures

#### Reset

| Name: | Reset |
| --- | --- |
| Identifier: | VIRCAM_img_cal_reset.tsf |
| Description: | Make Reset frame (reset-read only) with cold blank, |
| Parameters: | number of reset frames |
| Raw Frames: | FITS |
| Pipeline recipes: | vircam_reset_combine |

#### Dark

| Name: | Dark |
| --- | --- |
| Identifier: | VIRCAM_img_cal_dark.tsf |
| Description: | Make dark exposure (reset-read-read) with cold blank |
| Parameters: | integration time, number of integrations |
| Raw Frames: | FITS |
| Pipeline recipes: | vircam_dark_combine |

#### Dark Current

| Name: | Dark Current |
| --- | --- |
| Identifier: | VIRCAM_img_cal_darkcurrent.tsf |
| Description: | Make a series of dark exposures at a variety of different exposure times |
| Parameters: | List of integration times, and corresponding numbers of integrations |
| Raw Frames: | Sequence of FITS files |
| Pipeline recipes: | vircam_dark_combine |

#### Acquire Dome Screen

| Name: | Dome Screen |
| --- | --- |
| Identifier: | VIRCAM_img_acq_domescreen.tsf |
| Description: | Set instrument into IMAGING mode and select science filter. Move telescope to point at illuminated screen. |
| Parameters: | Filter, illumination combination |
| Raw Frames: | None |
| Pipeline recipes: | None |

#### Dome Flat

| Name: | Dome Flat |
| --- | --- |
| Identifier: | VIRCAM_img_cal_domeflat.tsf |
| Description: | Make a dome flat exposure (or sequence of exposures) suitable for calibrating IMAGING mode observations. |
| Parameters: | Filter, list of integration times and corresponding numbers of integrations |
| Raw Frames: | FITS files |
| Pipeline recipes: | vircam_dome_flat_combine |

#### Detector Linearity

| Name: | Linearity |
| --- | --- |
| Identifier: | VIRCAM_img_cal_linearity.tsf |
| Description: | Make series of dome flat exposures at a variety of exposure times. |
| Parameters: | Filter, List of integration times and corresponding numbers of integrations |
| Raw Frames: | FITS files |
| Pipeline recipes: | vircam_linearity_analyse |

#### Acquire Twilight Field

| Name: | Twilight |
| --- | --- |
| Identifier: | VIRCAM_img_acq_twilight.tsf |
| Description: | Select a dusk or dawn twilight field. Track (no autoguiding). Includes procedure to wait until sky brightness is appropriate. |
| Parameters: | Azimuth, Altitude, filter |
| Raw Frames: | None |
| Pipeline recipes: | None |

#### Twilight Flat

| Name: | Twilight Flat |
| --- | --- |
| Identifier: | VIRCAM_img_cal_twiflat.tsf |
| Description: | Take a series of exposures sufficient to make a twilight sky flat-field |
| Parameters: | List of integration times and corresponding numbers of integrations, or illumination level, depending on level of automation. Must give exposure time before repeating. |
| Raw Frames: | FITS files |
| Pipeline recipes: | vircam_twilight_combine |

#### Astrometric Calibration

No specific astrometric calibration templates are required as all science frames will be calibrated according to the procedure described in 6.2.2.

#### Photometric Calibration Standard Fields

All of the science modes 8.2.1.1 to 8.2.1.3 below may also be used to observe photometric standard fields for calibration.

### HOWFS mode exposures

Below IWS refers to the Instrument Workstation on which HOWFS processing is done.

#### HOWFS Dome Screen

| Name: | HOWFS Dome Screen |
| --- | --- |
| Identifier: | VIRCAM_howfs_acq_domescreen.tsf |
| Description: | Set camera into HOWFS mode and select HOWFS intermediate filter. Move telescope to dome illuminated screen, set tracking off and set illumination level. |
| Parameters: | Filter, screen illumination lamp combination. |
| Raw Frames: | None |
| IWS Procedures: | No |
| Pipeline recipes: | None |

#### HOWFS Dome Flat

| Name: | HOWFS Dome Flat |
| --- | --- |
| Identifier: | VIRCAM_howfs_cal_domeflat.tsf |
| Description: | Make a flat-field exposure (or exposures) suitable for calibrating HOWFS observations. |
| Parameters: | Filter, illumination combination, integration time, number of integrations |
| Raw Frames: | FITS |
| IWS Procedures: | Yes |
| Pipeline recipes: | None |

## Science Templates

### IMAGING mode exposures

The nesting of the observing loops is described in the same way as in the URD [AD2] using a shorthand based on the order of nesting of the loops for the 6 components, (F for filter, T for tile, P for pawprint, J for jitter, M for microstep, E for exposure), with the order of the letters indicating increasing nesting of the loop as one reads to the right.

#### Observe single Pawprint

##### Acquire

| Name: | Acquire |
| --- | --- |
| Identifier: | VIRCAM_img_acq_paw.tsf |
| Description: | Acquire single target. Check/Set camera to IMAGING mode, check/set camera position angle, check/select first science filter, all in parallel with a preset of telescope to new target, optionally (and usually) guide, optionally (and usually) activate LOWFS. i.e. nest Preset to tile ‘centre’ position Check/Set IMAGING mode in parallel Check/Set camera PA in parallel [default +X axis to +RA] Check/Set first filter in parallel For first pawprint Check/Offset telescope to pawprint position Acquire guide star LOWFS on two stars in parallel |
| Parameters: | Target coordinates, focal plane position to be at target position [e.g. centre of camera (default), or specified offset from centre of camera, or centre of a specified detector], camera position angle (E of N on sky, defaults to give +X to +RA), first filter, autoguiding required flag, if set (default) coordinates for 1 guide star from P2PP, LOWFS required flag, if set (default) 1 pair LOWFS stars found by P2PP. |
| Raw Frames: | None |
| Pipeline recipes: | None |

##### Observe_ (FJM)

| Name: | Observe_(FJM) |
| --- | --- |
| Identifier: | VIRCAM_img_obs_paw_fjm.tsf |
| Description: | The camera is already in IMAGING mode, position angle is set, first filter position is set, pointing at a target pawprint with autoguider and LOWFS on or off, (having been put in this state by the Acquire template). For each filter specified make a series of exposures at the positions specified by the patterns of jitter and microstep exposures, without changing guide star. The order of positions is determined to optimize speed. i.e. nest FJME For each demanded filter position (1 to F) Check/set demanded filter in For each jitter position (1 to J) Check/Move telescope (steps <30˝, same guide star) For each microstep (1 to M) Check/Move telescope (steps <3˝, same guide star) For each exposure (1 to E) Make exposure Next exposure Next microstep Next jitter Next filter |
| Parameters: | Number of filter positions F, and (if F>1) filter IDs; Number of jitter positions J, Number of microstep positions M nested at each jitter position; (if J >1) jitter pattern ID, jitter scale factor, and (if M=1) at each jitter position integration time, number of integrations; (if M>1) microstep pattern ID, microstep scale factor, and at each microstep position the integration time, number of integrations. |
| Raw Frames: | As many FITS files as there are exposures |
| Pipeline recipes: | vircam_microstep_interleave vircam_jitter_combine |
| Note: | The pipeline handles microstepped and jittered exposures differently. To just perform exposures at a fixed position set J=1 and M=1 To just perform a jitter pattern with no microsteps set M=1 To just perform a microstep pattern with no jitters set J=1 |

#### Observe set of Pawprints (e.g. to fill a tile)

##### Acquire Tile

| Name: | Acquire tile |
| --- | --- |
| Identifier: | VIRCAM_img_acq_tile.tsf |
| Description: | Acquire pawprint positions and parameters for a tile. The pattern of pawprint offsets is designed to fill in the gaps between the detectors to make an (almost) evenly sampled contiguous sky image – a tile. Check/Set camera to IMAGING mode, check/set camera position angle, check/select first science filter, obtain parameters specifying offsets required for a tiled exposure pattern, all in parallel with a preset of telescope to the “central null target” (defined to be the centre of the tile that will result), optionally (and usually) guide, optionally (and usually) activate LOWFS. This is similar to the Acquire Target template, but in addition obtains all the offsets (default 6) for the pawprints of the tile, and all the associated guide stars, and the jitter and microsteps at each pawprint position. |
| Parameters: | Coordinates of central null target, camera position angle (E of N on sky, defaults to give +X to +RA), filter list, tile ID (defaults to 2x3 giving offsets of 2 X steps of 95% detector width, 3 Y steps of 47.5% detector height), autoguiding required flag if set (default) list of guide stars found by P2PP (defaults to 6); LOWFS required flag, if set (default) list of LOWFS star pairs found by P2PP (defaults to 6 pairs). Keyword to show which of the Observe_( ) procedures is required. The selected Observe template parameters, LOWFS required flags, optional non-sidereal drift parameters. |
| Raw Frames: | None |
| Pipeline recipes: | None |

##### Observe_(PJM)

| Name: | Observe_(PJM) |
| --- | --- |
| Identifier: | VIRCAM_img_obs_tile_pjm.tsf |
| Description: | A template which cycles through all the pawprints of a tile (or set of offsets) to complete the tile (or set of offsets), with required jitters and microsteps at a fixed filter position. The camera is already in IMAGING mode, position angle is set, and filter position is set, pointing at first target pawprint with autoguider and LOWFS on or off, (having been put in this state by the Acquire Tile or Acquire Offsets template). Make a series of exposures at the positions specified by the pawprint positions and patterns of jitter and microstep exposures. The order of positions is determined to optimize speed. i.e. nest PJME For each pawprint position (1 to P) Check/offset telescope (steps 5-10΄) Acquire new guide and LOWFS stars For each jitter position (1 to J) Check/Move telescope (steps <30˝, same guide star) For each microstep (1 to M) Check/Move telescope (steps <3˝, same guide star) For each exposure (1 to E) Make exposure Next exposure Next microstep Next jitter Next pawprint |
| Parameters: | Are passed from VIRCAM_img_acq_tile.tsf or from VIRCAM_img_acq_offsets.tsf: Tile-pattern ID defining number and location of the pawprints [defaults to 6 locations], or list of offsets. Number of jitter positions J, Number of microstep positions M nested at each jitter position; (if J>1) jitter pattern ID, jitter scale factor, and (if M=1) at each jitter position integration time, number of integrations; (if M >1) microstep pattern ID, microstep scale factor, and at each microstep position the integration time, number of integrations. |
| Raw Frames: | As many FITS files as there are exposures |
| Pipeline Recipes: | vircam_microstep_interleave vircam_microstep_calibrate vircam_jitter_combine vircam_jitter_calibrate |
| Note | The pipeline handles microstepped and jittered exposures in a different way. |

##### Observe_(JPM)

| Name: | Observe_(JPM) |
| --- | --- |
| Identifier: | VIRCAM_img_obs_tile_jpm.tsf |
| Description: | At a fixed filter position, this template cycles through all the jitter positions (relative to the pawprint positions of a tile - or set of offsets), to complete the tile (or offsets), with required pawprints and microsteps e.g. at (0,0) jitter position observe all 6 pawprints, followed by at (10,10) jitter positions observe all 6 pawprints, etc. etc. This template has its jitter loop outside the pawprint loop, and is potentially useful if it is very important to minimize the variation of the final co-added point-spread function across the entire tile. However it increases the overheads by ~4 seconds per exposure due to the larger offsets and associated change of guide star with each offset. The camera is already in IMAGING mode, position angle is set, filter position is set, pointing at first target jitter position relative to pawprint position with autoguider and LOWFS on or off, (having been put in this state by the Acquire Tile or Acquire Offsets template). Make a series of exposures at the positions specified by the jitter positions, pawprint pattern and pattern of microstep exposures. i.e. nest JPME For each jitter position relative to each pawprint (1 to J) For each pawprint position (1 to P) Check/offset telescope (steps 5-10΄) to tile ‘centre’+ pawprint offset+ jitter position Acquire guide and LOWFS stars For each microstep (1 to M) Check/Move telescope (steps <3˝, same guide star) For each exposure (1 to E) Make exposure Next exposure Next microstep Next pawprint Next jitter |
| Parameters: | Are passed from either VIRCAM_img_acq_tile.tsf or from VIRCAM_img_acq_tile.tsf : Tile-pattern ID defining number and location of the pawprints [defaults to 6 locations], or list of offsets; Number of jitter positions J, Number of microstep positions M nested at each jitter position; (if J>1) jitter pattern ID, jitter scale factor, and (if M=1) at each jitter position integration time, number of integrations; (if M >1) microstep pattern ID, microstep scale factor, and at each microstep position the integration time, number of integrations. |
| Raw Frames: | As many FITS files as there are exposures |
| Pipeline Recipes: | vircam_microstep_interleave vircam_microstep_calibrate vircam_jitter_combine vircam_jitter_calibrate |
| Note | The pipeline handles microstepped and jittered exposures in a different way. |

##### Observe Tile

| Name: | Observe Tile |
| --- | --- |
| Identifier: | VIRCAM_img_obs_tile.tsf |
| Description: | Carry out the pattern of pawprint offsets to fill in the gaps between the detectors to make an (almost) evenly sampled contiguous sky image – a tile. The camera is already in IMAGING mode, position angle is set, first filter position is set, pointing at first target pawprint with autoguider and LOWFS on or off, (having been put in this state by the Acquire Tile template). Go to a series of pawprints at the offset pattern previously specified using the Acquire Tile template, changing guide star. At each pawprint execute, depending on a keyword, either the Observe_(FJM) procedure, or the Observe_(PJM) procedure. Or execute Observe_(JPM) procedure. The parameters of the selected Observe_( ) procedure must have been previously specified using the Acquire Tile template. For each pawprint (steps 5-10΄) (1 to P) Check/Offset telescope to pawprint central position Acquire guide star LOWFS on two stars in parallel either Observe_(FJM) or Observe_(PJM) Next pawprint Or Observe_(JPM) |
| Parameters: | Are passed from VIRCAM_img_acq_tile.tsf: Tile-pattern ID defining number and locations of the pawprints [defaults to 6 locations]; keyword to show which Observe_( ) procedure is required; all the parameters accepted by the Observe_( ) templates. |
| Raw Frames: | (Number of pawprint locations [6] x number of exposure in each pawprint) FITS files |
| Pipeline recipes: | as for pawprints as tile is not formed by pipeline |
| Note: | Pipeline produces pawprints, these are not tiled together. |

##### Acquire Offsets

| Name: | Acquire offsets |
| --- | --- |
| Identifier: | VIRCAM_img_acq_offsets.tsf |
| Description: | Acquire pawprint positions and parameters for any set of offset exposures. This is similar to the Acquire Tile template, but instead of acquiring the offsets from a selection of specific tile IDs they are provided explicitly to the template. This template is to allow versatility and can be used to produce offset patterns not contained within the set of Tile IDs, and to produce an arbitrary pattern. Check/Set camera to IMAGING mode, check/set camera position angle, check/select first science filter, obtain parameters specifying offsets required for an offset exposure pattern, all in parallel with a preset of telescope to the “central null target” (defined to be the centre of the region that will result), optionally (and usually) guide, optionally (and usually) activate LOWFS. |
| Parameters: | Coordinates of first target, camera position angle (E of N on sky, defaults to give +X to +RA), filter list, number of offsets P (defaults to 1), list of offset steps, autoguiding required flag (default set), if flag set list of guide stars (1 per offset position) found by P2PP, LOWFS required flag (default set), if set list of LOWFS star pairs (1 pair per position found by P2PP); keyword to show which Observe_( ) procedure is required; the selected Observe template parameters, LOWFS required flags, optional non-sidereal drift parameters. |
| Raw Frames: | None |
| Pipeline recipes: | None |

##### Observe Offsets

| Name: | Observe Offsets |
| --- | --- |
| Identifier: | VIRCAM_img_obs_offsets.tsf |
| Description: | Similar to Observe Tile except the offsets are not limited to a set of pre-defined offset patterns. The purpose is to allow the versatility of more general sets of offsets, rather than those offset pattern that have been predefined for produce a simple tile. Carry out the pattern of pawprint offsets to fill in the gaps between the detectors to make an (almost) evenly sampled contiguous sky image, i.e. a tile. The camera is already in IMAGING mode, position angle is set, first filter position is set, pointing at first target pawprint with autoguider and LOWFS on or off, (having been put in this state by the Acquire Offsets template). The list of offsets must have been defined in the Acquire Offsets template. Go to a series of pawprints at the P offsets previously specified using the Acquire Offsets template, changing guide star. At each pawprint execute, depending on a keyword, either the Observe_(FJM) procedure, or the Observe_(PJM) procedure. Or execute Observe_(JPM) procedure. The parameters of the selected Observe_( ) procedure must have been previously specified using the Acquire Offsets template. For each offset pawprint (unknown steps) (1 to P) Check/Offset telescope to pawprint central position Acquire guide star LOWFS on two stars in parallel either Observe_(FJM) or Observe_(PJM) Next pawprint Or Observe_(JPM) |
| Parameters: | Are passed from VIRCAM_img_acq_offsets.tsf: Number of offset locations and associated data; keyword to show which Observe_( ) procedure is required; all the parameters accepted by the Observe_( ) templates. |
| Raw Frames: | (Number of pawprint locations x number of exposure in each pawprint) FITS files |
| Pipeline recipes: | as for pawprints |
| Note | Pipeline produces pawprints, these are not merged. |

#### Observe set of Tiles

Three templates (FTPJME, TFPJME and TPFJME) that observe more than one tile were outlined in the URD [AD2]. However as explained there, no templates that observe multiple tiles have been defined for PDR. This is because the pipeline concentrates on individual tiles. The FTPJME nesting can nevertheless be achieved with multiple copies of the Observe Tile template within an OB – each specifying a different, single filter.

### HOWFS mode exposures

Below IWS refers to the Instrument Workstation on which HOWFS processing is done.

#### HOWFS Acquire

| Name: | HOWFS Acquire |
| --- | --- |
| Identifier: | VIRCAM_howfs_acq.tsf |
| Description: | Acquire a HOWFS (High-Order Wave Front Sensor) source. Set instrument into HOWFS mode which selects HOWFS intermediate filter. If guiding and LOWFS are required, set guide star and two LOWFS coordinate sets. |
| Parameters: | Target Coordinates, detector window, optionally guide and 2 LOWFS coordinates. |
| Raw frames: | None |
| IWS Procedures: | None |
| Pipeline recipes: | None |

#### HOWFS Wave front

| Name: | HOWFS wave front |
| --- | --- |
| Identifier: | VIRCAM_howfs_obs_wfront.tsf |
| Description: | Make a HOWFS wave front measurement for measuring the current residual from the active optics lookup table. This will typically be done only ~ twice per night, once at the start of the night, and once around midnight if necessary. |
| Parameters: | Filter, detector window, integration time, number of integrations |
| Raw Frames: | FITS |
| IWS Procedures: | Trigger HOWFS analysis system, forward coefficient residuals to TCS |
| Pipeline recipes: | None |

#### HOWFS Expose

| Name: | HOWFS Expose |
| --- | --- |
| Identifier: | VIRCAM_howfs_obs_exp.tsf |
| Description: | Make a HOWFS wave front measurement suitable for populating the active optics lookup tables in the TCS. This will be done only very occasionally [~quarterly] in engineering time and does not form part of the routine operations. |
| Parameters: | Filter, detector window, integration time, number of integrations |
| Raw Frames: | FITS |
| IWS Procedures: | Trigger HOWFS analysis system, produce look up table. |
| Pipeline recipes: | None |

## Instrument Health Templates

Instrument health monitoring templates are defined in [RD5] and are run on a regular basis. For example the instrument filter wheel is tested regularity for position repeatability, and this may determine how often to repeat a flat-field calibration with a particular science filter. The templates in [RD5] are not repeated here, since these monitoring outputs are not processed by the VISTA DFS and hence are not described in this Calibration Plan.

# Technical Programs

## TP-VIS1: Establishment of Secondary Standard Fields

This section outlines the procedures required to establish a network of secondary standard fields early in the operation of VIRCAM.

Figure 9-3 VISTA/WFCAM Standard Fields

| Name: | Secondary Standard Fields |
| --- | --- |
| Program Identifier: | TP-VIS1-IMA-PHO-0001 |
| Purpose: | Provide secondary standards for VISTA for routine calibrations (see Section 5) |
| Description: | A programme of observations around the primary standards is required to make direct measurements of all the secondary standards in the VIRCAM filter system. These observations will be repeated throughout the year to minimize the errors in the secondary star measurements, to identify variables, and to make full coverage in Right Ascension. The secondary standard fields selected are shown in Figure 9-3 which is a Hammer-Aitoff projection of targets selected from [RD8] and [RD9], and tabulated in Appendix A. For the equator, there are 63 fields with > 60 stars in one detector, with declination roughly in the range -10 to 10 degrees (Table A-2). (Restricting this to >100 stars per detector would restrict the RA coverage due to limiting fields to the galactic plane). Further fields are selected to be within 10° of the zeniths at VISTA (−24.67, Table A-1) and UΚIRT (+19.82, Table A-3). |
| Observing Conditions: | Photometric |
| Frequency: | Complete night at quarterly intervals over first 2 years of VIRCAM operations to ensure the photometric pedigree and accuracy of the standard fields |
| Special Conditions: | None |
| Analysis procedure: | A master catalogue of standard stars will be derived for each field with photometry in each of the VIRCAM filters. Photometry will be measured using standard VFDS pipeline procedures [RD1]. |
| Products: | Y, J, H, Κs magnitudes of ~1500 secondary standards in each field |
| Accuracies: | The target is 0.005 magnitude rms for secondary standards in each waveband after two years of repeated observations. |
| Responsible Person: | TBD |

# Format of Data Frames

## Principle

There is only one data format, used in both IMAGING and HOWFS modes. Data frames will be in ESO modified standard FITS format [RD10], the ESO modifications being limited to the hierarchical header proposal.   The headers are compliant with the final World Coordinate System (WCS) specification [RD11]. Data from the full set of chips is stored in Multi Extension Format (MEF) as 32-bit signed integers [RD10].  Offset 16-bit format is not used because data will be co-added in the data acquisition system before output. Though not a requirement, the integer format enables the use of highly efficient lossless compression.

## Model FITS header

A model FITS header for raw data is presented in Table 10-4. In addition to the header shown in the model, standard pipeline-processing keywords will be inserted into the data products. Table10-1 will be superseded by the formal issue of the VISTA DFS Data Interface Dictionary ([RD13] is in effect a preliminary version of this).

SIMPLE  =                    T / Standard FITS (NOST-100-2.0)

NAXIS   =                    0 / number of axes of data array

BITPIX  =                    8 / number of bits per pixel value

EXTEND  =                    T / FITS file extension may be present

RA      =           123.123457 / 00:00:00.123 RA of telescope

DEC     =           -12.123457 / -00:00:00.12 Dec of telescope

RADECSYS= 'ICRS    '           / Name of celestial reference frame

EQUINOX =               2000.0 / Equinox of celestial reference frame.

ORIGIN  = 'ESO     '           / European Southern Observatory

TELESCOP= 'ESO-VISTA'          / ESO <TEL>

INSTRUME= 'VIRCAM  '           / Instrument used

OBJECT  = 'Sirius  '           / Target description

IMAGETYP= 'OBJECT  '           / Exposure type

AIRMASS =              1.12346 / Averaged airmass

DATE    = '2004-03-29T13:37:57.000' / Date this file was written

DATE-OBS= '2004-12-25T09:00:00.123' / UTC date at start of exposure.

UTC     =            86399.123 / 00:00:00.123 UTC s at start since midnight

REQTIME =                5.000 / Requested integration time

EXPTIME =                5.123 / Actual integration time

LST     =            80000.123 / 00:00:00.123 LST seconds since midnight

MJD-OBS =       54321.12345678 / Modified Julian Date at start

OBSERVER= 'P Bunclarκ'         / Name of observer

PI-COI  = 'J Lewis '           / Name(s) of proposer(s)

COMMENT General comment

HISTORY Historical Fact

ESO-LOG

ORIGFILE= 'VIRCAM_Ima.1.fits'  / Original File Name

ARCFILE = 'VIRCAM.2006-03-05T07:25:0.000.fits      ' / Archive File name

CHECKSUM= 'Pd3jPc3hPc3h'       / ASCII 1s-complement checksum

RECIPE  = 'QUICΚ_LOOK'         / CASU Data-reduction recipe

OFFSTNUM=                 1234 / Value of first OBSNUM in current tile sequence

OBSNUM  =             12345678 / Observation Number

GRPNUM  =                  666 / Group number applied to all members

GRPMEM  =                    T / Group membership

STANDARD=                    F / Standard-star observation

NOFFSETS=                    6 / Number of offset positions in a field

OFFSET_I=                    2 / Serial Number of offset

NDITHER =                    6 / Number of positions in a tel dither pattern

DITHRNUM=                 1236 / Value of first OBSNUM in current dither sequenc

DITHER_I=                    3 / Serial number of this tel dither pattern

DITHER_X=                3.330 / X offset in dither pattern

DITHER_Y=                0.000 / Y offset in dither pattern

NUSTEP  =                    4 / Number of positions in microstep pattern

USTEPNUM=                 1237 / Value of first OBSNUM in current microstep sequ

USTEP_I =                    1 / Serial number of microstep pattern

USTEP_X =                1.123 / X offset in microstep pattern

USTEP_Y =                1.123 / Y offset in microstep pattern

HIERARCH ESO INS FILT1 ID       = 'J       '   / Filter identification device co

HIERARCH ESO INS FILT NAME1     = 'VIS#J/04_ESO843' / Filter Name

HIERARCH ESO INS FILT1 NO       =            2 / Filter wheel slot number

HIERARCH ESO INS FILT1 ENC      =         1234 / Filter wheel encoder position

HIERARCH ESO INS PRES1 ID       = 'P1      '   / ID of pressure sensor 1

HIERARCH ESO INS PRES1 NAME     = 'Cryostat'   / Location of pressure sensor 1

HIERARCH ESO INS PRES1 VAL      = 1.000000E-04 / Pressure sensor 1 reading

HIERARCH ESO INS PRES2 NAME     = 'Ambient '   / Location of pressure sensor 2

HIERARCH ESO INS PRES2 VAL      = 1.000100E+00 / Pressure sensor 2 reading

HIERARCH ESO INS PRES2 ID       = 'P2      '   / ID of pressure sensor 2

HIERARCH ESO INS FILT INPOS     =            T / filter in-position switch state

HIERARCH ESO INS FILT TIME      = '2006-03-05T16:17:18.123' / Last filt wheel in

HIERARCH ESO INS FILT ERROR     =           -6 / Last fil wheel index error

HIERARCH ESO INS TEMP29 TARGET  =       160.00 / Sensor 29 target

HIERARCH ESO INS TEMP30 TARGET  =       160.00 / Sensor 2 target

HIERARCH ESO INS TEMP1 ID       = 'ID1     '   / ID of sensor 1

HIERARCH ESO INS TEMP1 NAME     = 'Cryostat window cell ' / Location of sensor 1

HIERARCH ESO INS TEMP1 VAL      =       260.10 / Temperature sensor 1 reading

HIERARCH ESO INS TEMP2 ID       = 'ID2     '   / ID of sensor 2

HIERARCH ESO INS TEMP2 NAME     = 'Ambient air temperature ' / Location of senso

HIERARCH ESO INS TEMP2 VAL      =       260.10 / Temperature sensor 2 reading

HIERARCH ESO INS TEMP3 ID       = 'ID3     '   / ID of sensor 3

HIERARCH ESO INS TEMP3 NAME     = 'Baffle 7 '  / Location of sensor 3

HIERARCH ESO INS TEMP3 VAL      =       260.10 / Temperature sensor 3 reading

HIERARCH ESO INS TEMP4 ID       = 'ID4     '   / ID of sensor 4

HIERARCH ESO INS TEMP4 NAME     = 'Lens Barrel ' / Location of sensor 4

HIERARCH ESO INS TEMP4 VAL      =       260.10 / Temperature sensor 4 reading

HIERARCH ESO INS TEMP5 ID       = 'ID5     '   / ID of sensor 5

HIERARCH ESO INS TEMP5 NAME     = 'Liquid Nitrogen Tanκ ' / Location of sensor 5

HIERARCH ESO INS TEMP5 VAL      =       260.10 / Temperature sensor 5 reading

HIERARCH ESO INS TEMP6 ID       = 'ID6     '   / ID of sensor 6

HIERARCH ESO INS TEMP6 NAME     = 'Cold head 1 1st stage ' / Location of sensor

HIERARCH ESO INS TEMP6 VAL      =       260.10 / Temperature sensor 6 reading

HIERARCH ESO INS TEMP7 ID       = 'ID7     '   / ID of sensor 7

HIERARCH ESO INS TEMP7 NAME     = 'Cold head 1 2st stage ' / Location of sensor

HIERARCH ESO INS TEMP7 VAL      =       260.10 / Temperature sensor 7 reading

HIERARCH ESO INS TEMP8 ID       = 'ID8     '   / ID of sensor 8

HIERARCH ESO INS TEMP8 NAME     = 'Cold head 2 1st stage ' / Location of sensor

HIERARCH ESO INS TEMP8 VAL      =       260.10 / Temperature sensor 8 reading

HIERARCH ESO INS TEMP9 ID       = 'ID9     '   / ID of sensor 9

HIERARCH ESO INS TEMP9 NAME     = 'Cold head 2 2st stage ' / Location of sensor

HIERARCH ESO INS TEMP9 VAL      =       260.10 / Temperature sensor 9 reading

HIERARCH ESO INS TEMP10 ID      = 'ID10    '   / ID of sensor 10

HIERARCH ESO INS TEMP10 NAME    = 'Cold head 3 1st stage ' / Location of sensor

HIERARCH ESO INS TEMP10 VAL     =       260.10 / Temperature sensor 10 reading

HIERARCH ESO INS TEMP11 ID      = 'ID11    '   / ID of sensor 11

HIERARCH ESO INS TEMP11 NAME    = 'Cold head 3 2st stage ' / Location of sensor

HIERARCH ESO INS TEMP11 VAL     =       260.10 / Temperature sensor 11 reading

HIERARCH ESO INS TEMP12 ID      = 'ID12    '   / ID of sensor 12

HIERARCH ESO INS TEMP12 NAME    = 'Cryopump 1 ' / Location of sensor 12

HIERARCH ESO INS TEMP12 VAL     =       260.10 / Temperature sensor 12 reading

HIERARCH ESO INS TEMP13 ID      = 'ID13    '   / ID of sensor 13

HIERARCH ESO INS TEMP13 NAME    = 'Cryopump 2 ' / Location of sensor 13

HIERARCH ESO INS TEMP13 VAL     =       260.10 / Temperature sensor 13 reading

HIERARCH ESO INS TEMP14 ID      = 'ID14    '   / ID of sensor 14

HIERARCH ESO INS TEMP14 NAME    = 'WFS Plate ' / Location of sensor 14

HIERARCH ESO INS TEMP14 VAL     =       260.10 / Temperature sensor 14 reading

HIERARCH ESO INS TEMP15 ID      = 'ID15    '   / ID of sensor 15

HIERARCH ESO INS TEMP15 NAME    = 'WFS CCD 1 ' / Location of sensor 15

HIERARCH ESO INS TEMP15 VAL     =       260.10 / Temperature sensor 15 reading

HIERARCH ESO INS TEMP16 ID      = 'ID16    '   / ID of sensor 16

HIERARCH ESO INS TEMP16 NAME    = 'WFS CCD 2 ' / Location of sensor 16

HIERARCH ESO INS TEMP16 VAL     =       260.10 / Temperature sensor 16 reading

HIERARCH ESO INS TEMP17 ID      = 'ID17    '   / ID of sensor 17

HIERARCH ESO INS TEMP17 NAME    = 'WFS CCD 3 ' / Location of sensor 17

HIERARCH ESO INS TEMP17 VAL     =       260.10 / Temperature sensor 17 reading

HIERARCH ESO INS TEMP18 ID      = 'ID18    '   / ID of sensor 18

HIERARCH ESO INS TEMP18 NAME    = 'WFS CCD 4 ' / Location of sensor 18

HIERARCH ESO INS TEMP18 VAL     =       260.10 / Temperature sensor 18 reading

HIERARCH ESO INS TEMP19 ID      = 'ID19    '   / ID of sensor 19

HIERARCH ESO INS TEMP19 NAME    = 'Filter wheel shield ' / Location of sensor 19

HIERARCH ESO INS TEMP19 VAL     =       260.10 / Temperature sensor 19 reading

HIERARCH ESO INS TEMP20 ID      = 'ID20    '   / ID of sensor 20

HIERARCH ESO INS TEMP20 NAME    = 'Filter wheel hub ' / Location of sensor 20

HIERARCH ESO INS TEMP20 VAL     =       260.10 / Temperature sensor 20 reading

HIERARCH ESO INS TEMP21 ID      = 'ID21    '   / ID of sensor 21

HIERARCH ESO INS TEMP21 NAME    = 'Science detector 1 ' / Location of sensor 21

HIERARCH ESO INS TEMP21 VAL     =       260.10 / Temperature sensor 21 reading

HIERARCH ESO INS TEMP22 ID      = 'ID22    '   / ID of sensor 22

HIERARCH ESO INS TEMP22 NAME    = 'Science detector 2 ' / Location of sensor 22

HIERARCH ESO INS TEMP22 VAL     =       260.10 / Temperature sensor 22 reading

HIERARCH ESO INS TEMP23 ID      = 'ID23    '   / ID of sensor 23

HIERARCH ESO INS TEMP23 NAME    = 'Science detector 3 ' / Location of sensor 23

HIERARCH ESO INS TEMP23 VAL     =       260.10 / Temperature sensor 23 reading

HIERARCH ESO INS TEMP24 ID      = 'ID24    '   / ID of sensor 24

HIERARCH ESO INS TEMP24 NAME    = 'Science detector 4 ' / Location of sensor 24

HIERARCH ESO INS TEMP24 VAL     =       260.10 / Temperature sensor 24 reading

HIERARCH ESO INS TEMP25 ID      = 'ID25    '   / ID of sensor 25

HIERARCH ESO INS TEMP25 NAME    = 'Science detector 5 ' / Location of sensor 25

HIERARCH ESO INS TEMP25 VAL     =       260.10 / Temperature sensor 25 reading

HIERARCH ESO INS TEMP26 ID      = 'ID26    '   / ID of sensor 26

HIERARCH ESO INS TEMP26 NAME    = 'Science detector 6 ' / Location of sensor 26

HIERARCH ESO INS TEMP26 VAL     =       260.10 / Temperature sensor 26 reading

HIERARCH ESO INS TEMP27 ID      = 'ID27    '   / ID of sensor 27

HIERARCH ESO INS TEMP27 NAME    = 'Science detector 7 ' / Location of sensor 27

HIERARCH ESO INS TEMP27 VAL     =       260.10 / Temperature sensor 27 reading

HIERARCH ESO INS TEMP28 ID      = 'ID28    '   / ID of sensor 28

HIERARCH ESO INS TEMP28 NAME    = 'Science detector 8 ' / Location of sensor 28

HIERARCH ESO INS TEMP28 VAL     =       260.10 / Temperature sensor 28 reading

HIERARCH ESO INS TEMP29 ID      = 'ID29    '   / ID of sensor 29

HIERARCH ESO INS TEMP29 NAME    = 'FPA thermal plate (actual) ' / Location of se

HIERARCH ESO INS TEMP29 VAL     =       260.10 / Temperature sensor 29 reading

HIERARCH ESO INS TEMP30 ID      = 'ID30    '   / ID of sensor 30

HIERARCH ESO INS TEMP30 NAME    = 'Cryostat tube outside (actual) ' / Location o

HIERARCH ESO INS TEMP30 VAL     =       260.10 / Temperature sensor 30 reading

HIERARCH ESO INS TEMP31 ID      = 'ID31    '   / ID of sensor 31

HIERARCH ESO INS TEMP31 NAME    = 'undecided temperature sensor ' / Location of

HIERARCH ESO INS TEMP31 VAL     =       260.10 / Temperature sensor 31 reading

HIERARCH ESO INS TEMP32 ID      = 'ID32    '   / ID of sensor 32

HIERARCH ESO INS TEMP32 NAME    = 'undecided temperature sensor ' / Location of

HIERARCH ESO INS TEMP32 VAL     =       260.10 / Temperature sensor 32 reading

HIERARCH ESO INS HOWFS DATE     = '2006-03-05T01:02:03.123' / Time of new coefs

HIERARCH ESO OBS DID         = 'ESO-VLT-DIC.OBS' / OBS Dictionary

HIERARCH ESO OBS ID          =              666 / Observation block ID

HIERARCH ESO OBS NAME        = 'DITHER                          ' / OB name

HIERARCH ESO OBS GRP         = 'ABCD    '      / linκed blocks

HIERARCH ESO OBS OBSERVER    = 'Bunclarκ                        ' / Observer Nam

HIERARCH ESO OBS PI-COI ID   =              162 / ESO internal PI-COI ID

HIERARCH ESO OBS PI-COI NAME = 'Lewis                           ' / PI-COI name

HIERARCH ESO OBS PROG ID     = '68.A-0281(A)        ' / ESO program identificati

HIERARCH ESO OBS TPLNO       =                2 / Template number within OB

HIERARCH ESO OBS TARG NAME   = 'South Pole                      ' / OB target na

HIERARCH ESO OBS START       = '2006-03-05T07:20:00.123' / OB start time

HIERARCH ESO OBS EXECTIME    =                0 / Expected execution time

HIERARCH ESO TPL PRESEQ      = 'VIRCAM_img_obs_Dither.seq' / Sequencer script

HIERARCH ESO TPL START       = '2006-03-05T07:20:00.123' / TPL start time

HIERARCH ESO TPL DID         = 'ESO-VLT-DIC.TPL-1.9           ' / Data dictionar

HIERARCH ESO TPL ID          = 'VIRCAM_img_obs_Dither                          '

HIERARCH ESO TPL NAME        = 'VIRCAM Dithered mode sequence                  '

HIERARCH ESO TPL NEXP        =                6 / Number of exposures within tem

HIERARCH ESO TPL EXPNO       =                2 / Exposure number within templat

HIERARCH ESO TPL VERSION     = '@(#) $Revision: 1.5 $' / Version of the template

HIERARCH ESO DPR DID         = 'ESO-VLT-DIC.DPR-1.8' / DPR Dictionary

HIERARCH ESO DPR CATG        = 'SCIENCE                       ' / Observation ca

HIERARCH ESO DPR TYPE        = 'OBJECT                        ' / Observation ty

HIERARCH ESO DPR TECH        = 'IMAGE                         ' / Observation te

HIERARCH ESO TEL FOCU LEN       =        4.120 / Focal length (m)

HIERARCH ESO TEL FOCU SCALE     =       24.000 / Focal scale (arcsec/mm)

HIERARCH ESO TEL FOCU VALUE     =    12345.120 / M2 setting (mm)

HIERARCH ESO TEL PARANG END     =       45.000 / Parallactic angle at end (deg)

HIERARCH ESO TEL PARANG START   =       47.000 / Parallactic angle at start (deg

HIERARCH ESO TEL AIRM END       =        1.001 / Airmass at end

HIERARCH ESO TEL AIRM START     =        1.002 / Airmass at start

HIERARCH ESO TEL TRAΚ STATUS    = 'ON        ' / Tracking status

HIERARCH ESO TEL TRAΚ RATEA     =     0.000000 / Tracking rate in RA (arcsec/sec

HIERARCH ESO TEL TRAΚ RATED     =     0.000000 / Tracking rate in DEC (arcsec/se

HIERARCH ESO TEL DOME STATUS    = 'FULLY-OPEN' / Dome status

HIERARCH ESO ADA GUID STATUS    = 'ON        ' / Status of autoguider

HIERARCH ESO ADA GUID RA        =   180.000000 / 00:00:00.123 Guide star RA J200

HIERARCH ESO ADA GUID DEC       =    -45.00000 / %DEGREE Guide star DEC J2000

HIERARCH ESO ADA POSANG         =     33.00000 / Position angle at start

HIERARCH ESO ADA ABSROT START   =      2.00000 / Abs rot angle at exp start (deg

HIERARCH ESO TEL ID             = 'v 3.45                        ' / TCS version

HIERARCH ESO TEL DID            = 'ESO-VLT-DIC.TCS-1.33          ' / Data dictio

HIERARCH ESO TEL DATE           = '2006-05-03' / TCS installation date

HIERARCH ESO ADA ABSROT END     =      3.00000 / Abs rot angle at exp end (deg)

HIERARCH ESO ADA ABSROT PPOS    = posit        / sign of probe position

HIERARCH ESO ADA WFS1 RA        =    12.123457 / RA of WFS star 1

HIERARCH ESO ADA WFS1 DEC       =   -75.987654 / Dec of WFS star 1

HIERARCH ESO ADA WFS2 RA        =    12.123457 / RA of WFS star 2

HIERARCH ESO ADA WFS2 DEC       =   -75.987654 / Dec of WFS star 2

HIERARCH ESO TEL ALT            =       80.000 / Alt angle at start (deg)

HIERARCH ESO TEL AZ             =       10.000 / Az angle at start (deg) S=0,W=9

HIERARCH ESO TEL GEOELEV        =         2335 / Elevation above sea level (m)

HIERARCH ESO TEL GEOLAT         =     -29.2543 / Tel geo latitute (+=North) (deg

HIERARCH ESO TEL GEOLON         =     -70.7346 / Tel geo longitude (+=East) (deg

HIERARCH ESO TEL OPER           = 'Senor Operador                ' / Telescope O

HIERARCH ESO TEL FOCU ID        = 'CA        ' / Telescope focus station ID

HIERARCH ESO TEL TH M1 TEMP     =         8.12 / M1 superficial temperature

HIERARCH ESO TEL MOON RA        =    10.000000 / 00:00:00.123 RA (J2000) (deg)

HIERARCH ESO TEL AMBI FWHM START=         0.50 / Observatory Seeing queried from

HIERARCH ESO TEL AMBI FWHM END  =         0.70 / Observatory Seeing queried from

HIERARCH ESO TEL AMBI TEMP      =         4.20 / Observatory ambient temperature

HIERARCH ESO TEL AMBI RHUM      =            5 / Observatory ambient relative hu

HIERARCH ESO TEL AMBI WINDDIR   =          340 / Observatory ambient wind direct

HIERARCH ESO TEL AMBI WINDSP    =        15.00 / Observatory ambient wind speed

HIERARCH ESO TEL MOON DEC       =     20.00000 / %DEGREE DEC (J2000) (deg)

HIERARCH ESO TEL ENCL FFLAMP1 ID= '123     '   / Dim tungsten lamp pair

HIERARCH ESO TEL ENCL FFLAMP1 NAME= 'VIS_DOM_DIM' / Dim tungsten lamp pair

HIERARCH ESO TEL ENCL FFLAMP1 STATE= 'OFF     ' / ON/OFF state of flat lamp 1

HIERARCH ESO TEL ENCL FFLAMP2 ID= '234     '   / Bright tungsten lamp pair

HIERARCH ESO TEL ENCL FFLAMP2 NAME= 'VIS_DOM_BRIGHT' / Bright tungsten lamp pair

HIERARCH ESO TEL ENCL FFLAMP2 STATE= 'ON      ' / ON/OFF state of flat lamp 2

HIERARCH ESO TEL ENCL FFLAMP3 ID= '345     '   / Halogen lamp pair

HIERARCH ESO TEL ENCL FFLAMP3 NAME= 'VIS_DOM_HALOGEN' / Dim tungsten lamp pair

HIERARCH ESO TEL ENCL FFLAMP3 STATE= 'OFF     ' / ON/OFF state of flat lamp 3

HIERARCH ESO TEL ENCL MOONSCR STEP=            1 / Moonscreen positions step

HIERARCH ESO TEL ENCL WINDSCR1 STATE= 'OPEN    ' / Louvre state

HIERARCH ESO TEL ENCL WINDSCR2 STATE= 'UP      ' / up/down slide state

HIERARCH ESO TEL ENCL VENT1 STATE= 'SHUT    '  / Vent 1 door state

HIERARCH ESO TEL ENCL VENT2 STATE= 'HALF    '  / Vent 2 door state

HIERARCH ESO TEL ENCL VENT3 STATE= 'OPEN    '  / Vent 3 door state

HIERARCH ESO TEL M2 LOOP1 STATE = 'CLOSED  '   / Focus-loop switch state

HIERARCH ESO TEL M2 LOOP2 STATE = 'OPEN    '   / Centroiding-loop switch state

HIERARCH ESO TEL M2 LOOP3 STATE = 'CLOSED  '   / Tilt-loop switch state

HIERARCH ESO TEL M2 LOOP4 STATE = 'CLOSED  '   / Astigmatic-loop switch state

HIERARCH ESO TEL M2 LOOP5 STATE = 'CLOSED  '   / Trefoil-loop switch state

HIERARCH ESO TEL M2 CENX        = 1.510000     / X-Centre reading 1

HIERARCH ESO TEL M2 CENY        = 1.520000     / Y-Centre reading 2

HIERARCH ESO TEL M2 TILTX       = 1.530000     / X-tilt reading 3

HIERARCH ESO TEL M2 TILTY       = 1.540000     / Y-tilt reading 4

HIERARCH ESO TEL M1 ACTUATORFAILED=            1 / Number of failed actuator

END

XTENSION= 'IMAGE   '           / Extension first keyword

NAXIS   =                    2 / number of axes of data array

NAXIS1  =                 2048 / Size of first axis

NAXIS2  =                 2048 / Size of second axis

BITPIX  =                   32 / number of bits per pixel value

EXTNAME = 'WIN1.CHIP1.OUT1'    / FITS extension name

EXTVER  =                    1 / Detector index

INHERIT =                    T / Extension inherits primary header

DET_LIVE=                    T / This detector is alive

RADECSYS= 'ICRS    '           / Name of celestial reference frame

EQUINOX =               2000.0 / Equinox of celestial reference frame.

CTYPE1  = 'RA---ZPN'           / Type of celestial axis 1

CTYPE2  = 'DEC--ZPN'           / Type of celestial axis 2

CRPIX1  =              6860.80 / Pixel coordinate of reference in axis 1

CRPIX2  =             -3507.20 / Pixel coordinate of reference in axis 2

CRVAL1  =     270.000000000000 / RA of reference point

CRVAL2  =     -75.000000000000 / Dec of reference point

CD1_1   = -9.444444E-05        / Transformation matrix element

CD1_2   = 0.000000E+00         / Transformation matrix element

CD2_1   = 0.000000E+00         / Transformation matrix element

CD2_2   = 9.444444E-05         / Transformation matrix element

PV2_1   = 1.000000E+00         / linear term in ZPN

PV2_2   = 0.000000E+00         / quadratic term in ZPN

PV2_3   = 4.200000E+01         / cubic term in ZPN

PV2_4   = 0.000000E+00         / forth-order term in ZPN

PV2_5   = 0.000000E+00         / fifth-order term in ZPN

HIERARCH ESO DET MINDIT      =        1.2345678 / Minimum DIT

HIERARCH ESO DET MODE NAME   = 'UNKNOWN '      / DCS Detector Mode

HIERARCH ESO DET MODEi NAME  = 'UNKNOWN '      / DCS Detector Mode

HIERARCH ESO DET NC IDIV     =                0 / # of Sub-Divs for R-O

HIERARCH ESO DET NC NSAMPPIX =                0 / # of Samples/Pixel

HIERARCH ESO DET NCORRS      =                0 / Read-Out Mode

HIERARCH ESO DET RSPEED      =                0 / Read-Speed Factor

HIERARCH ESO DET IRACE ADCi NAME= 'UNKNOWN '   / Name for ADC Board

HIERARCH ESO DET IRACE ADCi HEADER=                0 / Header of ADC Board

HIERARCH ESO DET IRACE ADCi ENABLE=                0 / Enable ADC Board (0/1)

HIERARCH ESO DET IRACE ADCi FILTERi=                0 / ADC Filter Adjustment

HIERARCH ESO DET IRACE ADCi DELAY=                0 / ADC Delay Adjustment

HIERARCH ESO DET NCORRS NAME = 'UNKNOWN '      / Read-Out Mode Name

HIERARCH ESO DET NDIT        =                0 / # of Sub-Integrations

HIERARCH ESO DET NDITSΚIP    =                0 / DITs sκipped at 1st.INT

HIERARCH ESO DET NDSAMPLES   =                0 / # of Non-Dest. Samples

HIERARCH ESO DET NDSΚIP      =                0 / Samples sκipped per DIT

HIERARCH ESO DET RSPEEDADD   =                0 / Read-Speed Add

HIERARCH ESO DET VOLTi CLΚHINMi= 'UNKNOWN '    / Name of High Clocκ

HIERARCH ESO DET VOLTi CLΚHITi=           1.2346 / Tel Value High Clocκ

HIERARCH ESO DET VOLTi CLΚHIi=           1.2346 / Set Value High Clocκ

HIERARCH ESO DET VOLTi CLΚLONMi= 'UNKNOWN '    / Name of Low Clocκ

HIERARCH ESO DET IRACE SEQCONT=          UNKNOWN / Sequencer Cont. Mode

HIERARCH ESO DET IRACE SEQINT=          UNKNOWN / Sequencer Intr. at Stop

HIERARCH ESO DET VOLTi CLΚLOTi=           1.2346 / Tel Value Low Clocκ

HIERARCH ESO DET VOLTi DCNMi = 'UNKNOWN '      / Name of DC Voltage

HIERARCH ESO DET VOLTi DCTAi =           1.2346 / Tel Value 1 for DC

HIERARCH ESO DET VOLTi CLΚLOi=           1.2346 / Set Value Low Clocκ

HIERARCH ESO DET VOLTi DCi   =           1.2346 / Set Value DC Voltage

HIERARCH ESO DET CHIP TYPE   = 'UNKNOWN '      / The Type of Det Chip

HIERARCH ESO DET CON OPMODE  = 'UNKNOWN '      / Operational Mode

HIERARCH ESO DET FRAM TYPE   = 'UNKNOWN '      / Frame type

HIERARCH ESO DET CHOP CYCSΚIP=                0 / # of Chop Cycles to Sκip

HIERARCH ESO DET CHOP NCYCLES=                0 / # of Chop Cycles

HIERARCH ESO DET CHOP ST     =          UNKNOWN / Chopping On/Off

HIERARCH ESO DET CHOP FREQ   = 0.000000        / Chopping Frequency

HIERARCH ESO DET CHIP ID     = 'UNKNOWN '      / Detector ID

HIERARCH ESO DET CHIP NAME   = 'UNKNOWN '      / Detector name

HIERARCH ESO DET CHIP NX     =                0 / Pixels in X

HIERARCH ESO DET CHIP NY     =                0 / Pixels in Y

HIERARCH ESO DET CHIP PXSPACE= 0.000e+00       / Pixel-Pixel Spacing

HIERARCH ESO DET EXP NO      =                0 / Exposure Number

HIERARCH ESO DET FRAM UTC    = 'UNKNOWN '      / Time Recv Frame

HIERARCH ESO DET FRAM NO     =                0 / Frame number

HIERARCH ESO DET VOLTi DCTBi =           1.2346 / Tel Value 2 for DC

HIERARCH ESO DET WIN NX      =                0 / # of Pixels in X

HIERARCH ESO DET DID         = 'UNKNOWN '      / Dictionary Name and Revision

HIERARCH ESO DET DIT         =        1.2345678 / Integration Time

HIERARCH ESO DET DITDELAY    =            1.235 / Pause Between DITs

HIERARCH ESO DET EXP UTC     = 'UNKNOWN '      / File Creation Time

HIERARCH ESO DET EXP NAME    = 'UNKNOWN '      / Exposure Name

HIERARCH ESO DET WIN NY      =                0 / # of Pixels in Y

HIERARCH ESO DET WIN STARTX  = 0.000000        / Lower Left X Ref

HIERARCH ESO DET WIN STARTY  = 0.000000        / Lower left Y Ref

HIERARCH ESO DET WIN TYPE    =                0 / Win-Type

END

The section between the two ENDs repeating as appropriate for the next 15 extensions.

Table 10-4 FITS Example Header

# All-Sky IR Survey Fields

Name			RA		DEC		Num

9103         00 33 15.2  -39 24 10        59

9106         03 26 53.9  -39 50 38        65

9115         05 36 44.8  -34 46 39       133

HD38921      05 47 22.19 -38 13 51.3     125

9123         06 59 45.6  -30 13 44       530

9132         08 25 36.1  -39 05 59      1434

9133         08 27 12.5  -25 08 01       577

9137         09 15 50.5  -36 32 34       639

FS140        17 13 22.65 -18 53 33.8    1061

HD161743     17 48 57.93 -38 07 07.5    3223

FS34         20 42 34.73 -20 04 34.8     221

Table A-1 Southern Standards

Name			RA		DEC		Num

HD1160       00 15 57.30 +04 15 04.0      82

BRI0021      00 24 24.60 -01 58 22.0      72

FS2          00 55 09.93 +00 43 13.1      78

FS3          01 04 21.63 +04 13 36.0      79

FS105        01 19 08.19 +07 34 11.5      79

FS6          02 30 16.64 +05 15 51.1      82

T832-38078   03 04 02.00 +00 45 52.0      94

FS110        03 41 02.22 +06 56 15.9      73

FS10         03 48 50.20 -00 58 31.2      74

FS11         04 52 58.92 -00 14 41.6     128

FS119        05 02 57.44 -01 46 42.6     173

SAO112626    05 19 17.16 +01 42 16.1     176

S840-F       05 42 32.10 +00 09 04.0     293

FS13         05 57 07.59 +00 01 11.4     375

HD40335      05 58 13.52 +01 51 23.0     342

9118         06 22 43.7  -00 36 30       593

S842-E       06 22 43.70 -00 36 30.0     593

SA98-653     06 52 04.95 -00 18 18.3     958

FS121        06 59 46.82 -04 54 33.2     928

FS14         07 24 14.40 -00 33 04.1     507

RU149D       07 24 15.36 -00 32 47.9     514

P545-C       08 29 25.10 +05 56 08.0     158

LHS2026      08 32 30.50 -01 34 37.0     162

S705-D       08 36 12.50 -10 13 39.0     262

FS18         08 53 35.51 -00 36 41.7     121

FS124        08 54 12.60 -08 05 03.0     173

HD77281      09 01 38.01 -01 28 34.8     122

GL347A       09 28 53.50 -07 22 15.0     122

S708-D       09 48 56.40 -10 30 32.0     108

P550-C       10 33 51.80 +04 49 05.0      71

FS20         11 07 59.93 -05 09 26.1      79

HD121968     13 58 51.17 -02 54 52.3      65

S791-C       13 17 29.60 -05 32 37.0      65

HD129655     14 43 46.44 -02 30 20.0      93

FS136        14 59 32.05 -00 06 17.0      74

T868-53850   15 00 26.40 -00 39 29.0      81

T868-110639  15 10 17.00 -02 41 05.0     100

PG1528+062B  15 30 39.00 +06 01 13.0     119

S870-T       15 39 03.50 +00 14 54.0     135

FS137        16 26 42.72 +05 52 20.3     170

SA108-475    16 37 00.00 -00 34 39.0     241

9170         17 27 22.2  -00 19 25       394

SA109-381    17 44 12.00 -00 20 33.0     574

HD161903     17 48 19.22 -01 48 29.7     811

FS35         18 27 13.52 +04 03 09.4    1502

FS143        18 29 53.79 +01 13 29.9     787

FS144        18 29 56.90 +01 12 47.1     814

SA110-232    18 40 52.00 +00 01 55.0    3177

L547         18 51 15.60 -04 16 02.0    3773

S808-C       19 01 55.40 -04 29 12.0    3043

GL748        19 12 14.60 +02 53 11.1    3007

FS148        19 41 23.52 -03 50 56.1     856

S813-D       20 41 05.10 -05 03 43.0     261

SA112-822    20 42 54.00 +00 15 02.0     289

P576-F       20 52 47.30 +06 40 05       288

GL811.1      20 56 46.60 -10 26 54.6     178

HD201941     21 12 45.32 +02 38 33.9     226

FS29         21 52 25.36 +02 23 20.7     138

9185         22 02 05.7  -01 06 02       114

FS30         22 41 44.72 +01 12 36.5     104

FS31         23 12 21.60 +10 47 04.1      80

FS32         23 16 12.37 -01 50 34.6      85

FS154        23 18 10.08 +00 32 55.6      85

Table A-2 Equatorial Standards

Name			RA		DEC		Num

HD3029       00 33 39.53 +20 26 01.7      58

FS109        03 13 24.16 +18 49 38.4     103

FS115        04 23 18.17 +26 41 16.4     182

FS12         05 52 27.66 +15 53 14.3     689

FS120        06 14 01.44 +15 09 58.3     774

P309-U       07 30 34.50 +29 51 12.0     187

FS21         11 37 05.15 +29 47 58.4      52

FS23         13 41 43.57 +28 29 49.5     283

HD136754     15 21 34.53 +24 20 36.1      99

FS141        17 48 58.87 +23 17 43.7     245

FS149        20 00 39.25 +29 58 40.0    3317

FS152        22 27 16.12 +19 16 59.2     146

Table A-3 Northern Standards

_oo_