## Sheet1
IR CAMERA
[6 empty] | Document Title: |  | IR Detectors Performance and Status
[6 empty] | Document Number: |  | VIS-LST-ATC-06032-1000
[6 empty] | Issue: |  | Draft 0.3
[6 empty] | Date: |  | 23 November 2004
|  |  |  | Document |  | Naidu Bezawada |  |  | Signature
|  |  |  | Prepared by: |  | Electronics Engineer |  |  | and Date
|  |  |  | Document |  | Mel Strachen |  |  | Signature
|  |  |  | Approved by: |  | ATC Project Manager |  |  | and Date
|  |  |  | Document |  | Gavin Dalton |  |  | Signature
|  |  |  | Reviewed by: |  | IR Camera Project Scientist |  |  | and Date
|  |  |  | Document |  | Martin Caldwell |  |  | Signature
|  |  |  | Reviewed by: |  | IR Camera Systems Engineer |  |  | and Date
|  |  |  | Document |  | Kim Ward |  |  | Signature
|  |  |  | Released by: |  | IR Camera Project Manager |  |  | and Date
The information contained in this document is strictly confidential and is intended for the addressee only.
The unauthorised use, disclosure, copying, alteration or distribution of this document is strictly prohibited and may be unlawful
[12 empty] | Doc Number: | VIS-LST-ATC-06032-1000
[5 empty] | IR Detectors | [6 empty] | Date: | 23 November 2004
[5 empty] | Performance and Status | [6 empty] | Issue: | Draft 0.3
IR Camera | [11 empty] | Author: | Naidu Bezawada
Module No. | Grade | Status | QE (%) | Non-unif. | Dark (e/s) | Noise (e rms) | Full well(ke) | Non-lin. (%) | Conv. Gain | Mean DC | Persistence | Operability | Comments
|  |  | J / H / K |  | 0.5V / 0.7V | Temp / | 0.5V / 0.7V | 0.5V / 0.7V |  |  |  | (%)
Pix - to -pix
Process - Thin substrate
Module - 15 | Eng. | Tests completed | 59.1 |  | 1.44 | 22 | 130 | 1.9 | 2.67 | 3.192 |  |  | Engeneering grade
|  |  | 56.6 | [9 empty] | Large area defect
|  |  | 48.8 | [9 empty] | Not to publish data
Imperial hardware
Module - 16 | Eng. | Tests completed | 76 |  | 2.3 | 19.5 | 105 | 1.3 | 3.19 | 3.2 |  |  | Good engineering grade
|  |  | 76 | [9 empty] | detector
|  |  | 65 | [9 empty] | Imperial hardware
15 - 17% high dark pixels
Module - 18 | Eng. | Tests completed | 92.7 |  | 2.37 | 22.7 | 115 | 1.4 | 2.91 | 3.179 |  | 93.15 | Good engineering grade
|  |  | 96.5 |  |  | 25.1 | [6 empty] | detector, 40% high dark
|  |  | 96 | [9 empty] | pixels, Imperial hardware
Process - Thick substrate
Module - 22 | Sci. | Tests completed | 71 | 7.4 | 1.7 | 17.4 | 102 | 3.2 | 3.58 | 3.231 | <32e in ~45s | 99.18 | Science grade
|  | Accepted | 74 | 7.8 | 2 | 18.8 | 156 | 5.1 | 3.75 |  |  |  | AR coating - old process
|  |  | 75 | 5.9 | [8 empty] | Flatness ??
Module - 23 | Sci. | Tests completed | 71.8 | 5.9 | 1.28 | 17.2 | 96 | 3.3 | 3.81 | 3.237 | 17 | 98.68 | Science grade
|  | Accepted | 73.3 | 5.7 |  | 22.5 | 132 | [5 empty] | AR coating - old process
|  |  | 71.6 | 5.6 | [8 empty] | Flatness = 16µm p-v
Module - 24 |  | ROIC failed
Returned to RVS
Module - 25 | Sci. | Tests completed | 86 | 5 |  | 17.9 | 90 | 2.24 | 3.68 | 3.239 | 24.7 | 98.55 | Science grade
|  | Accepted | 89.4 | 5.2 | 0.86 | 28.7 | 124 | 3.25 |  |  |  |  | High QE
|  |  | 86.7 | 4.4 | [8 empty] | Robust AR coating
Flatness = 24µm p-v
Module - 26 | Eng. | Tests completed | 93.2 | 8.3 | 2.3 | 16.3 | 84 | 3.8 | 3.69 | 3.245 | 37.3 | 58.1 | Engineering grade
|  | Under evaluation | 96.5 | 6.9 | 2.38 | 16.7 | 115 | 5.6 | 4 | 3.261 |  |  | Flatness = 24µm p-v
|  |  | 91.5 | 7 | [8 empty] | Limited by high dark
~40% hot pixels
Module - 27 |  | Cracked
Not shipped
Sold to IFA
Module - 28 | Eng. | Tests completed | 64.1 | 11.9 |  | 28.6 | 98 | 3.6 | 4.08 | 3.266 | 26.2 | 96.28 | Engineering grade
|  | Under evaluation | 67.8 | 11.1 | 0.89 | 58.4 | 135 | 4.6 |  |  |  |  | Flatness = 18µm p-v
|  |  | 67.1 | 10.6 | [8 empty] | Numerous defects in the
centre, high read nosie
Module - 29 |  | ROIC failed
Not shipped
Module - 30 | Sci. | Tests completed | 93.5 | 6.6 | 3.85 | 22.2 | 146 | 9 | 2.32 | 3.148 | 13.6 | 93.15 | Science grade
|  | Accepted | 96.7 | 6.9 | 4 | 27.2 | 185 | 12.4 |  |  |  |  | Large hot pixels
|  |  | 92.6 | 5.9 | [8 empty] | High non-linearity
Dark current relaxed to
accept as science grade
Module - 31 | Sci. | Tests Completed | 94.2 | 5.4 | 0.97 | 15.5 | 93 | 5.7 | 3.44 | 3.229 | 9.6 | 98.3 | Science grade
|  | Accepted | 97.7 | 4.9 | 1.07 | 19.1 | 130 | 3.7 | 3.68 | 3.241 |  |  | High QE, good uniformity
|  |  | 96.5 | 5.1 | [8 empty] | Low noise, high operability
Module - 32 | Eng. | Under squeezed
Not shipped
Module - 33 | Sci. | Tests Completed | 95.5 | 5.4 | 1.13 | 22.7 | 97 | 5.9 | 3.18 | 3.202 | 14.2 | 98.15 | Science grade
|  | Accepted | 98.7 | 6 | 1.16 | 25.1 | 127 | 7.1 | 3.51 | 3.217 |  |  | High QE, good uniformity
|  |  | 93.8 | 5.2 | [8 empty] | high operability
Module - 34 |  | ROIC failed
Not shipped
Module - 35 | Sci | Tests Completed | 92.4 | 4.9 | 0.8 | 18.7 | 87 | 3.5 | 4.39 | 3.285 | 7.8 | 97.3 | Science grade
|  | Under evaluation | 96.3 | 4.8 | 0.81 | 21.4 | 120 | 5.2 | 4.69 | 3.299 |  |  | Except for three large area
|  |  | 90.7 | 4.3 | [8 empty] | defects ~1.6% of pixels in total
High QE, good uniformity,
Low persistence
Module - 36 | Sci | Under tests
Module - 37 | Sci. | Awaiting tests
Module - 38 | Sci. | Awaiting tests
Module - 39 | Sci | Expected
Module - 40 | Sci | Expected | [10 empty] | High dark
Waiver reqired
Module - 41 | Sci | Expected
Module - 42 | Sci | Expected
Module - 43 | Sci | Expected | [10 empty] | High dark
Waiver required
Module - 44 |  | Expected