###### IR CAMERA

##### Document Title: 	 VIRGO Module SCA-35 Test Report

Document Number: 	 VIS-TRE-ATC-06032-1035

Issue: 	 Draft 0.1

Date: 	 19 November 2004

| Document Prepared By: | Naidu Bezawada Electronics Engineer | Signature and Date: |  |
| --- | --- | --- | --- |
| Document Approved By: | Mel Strachan ATC Project Manager | Signature and Date: |  |
| Document Reviewed By: | Gavin Dalton IR Camera Project Scientist | Signature And Date: |  |
| Document Reviewed By: | Martin Caldwell IR Camera Systems Engineer | Signature and Date: |  |
| Document Released By: | Kim Ward IR Camera Project Manager | Signatures and Date: |  |

The information contained in this document is strictly confidential and is intended for the addressee only.  The unauthorised use, disclosure, copying, alteration or distribution of this document is strictly prohibited and may be unlawful.

# Change Record

| Issue | Date | Section(s) Affected | Description of Change/Change Request Reference/Remarks |
| --- | --- | --- | --- |
| Draft 0.1 | 19/11/04 | All | New document |

TABLE OF CONTENTS

Change Record	2

1.	PERFORMANCE SUMMARY	4

2.	DC Uniformity	5

3.	System gain and conversion gain	7

4.	Read noise	12

5.	Dark generation	15

6.	Quantum Efficiency	17

7.	Persistence	23

8.	Inoperable Pixels	26

9.	Temporal Stability	27

10.	Cosmetics	29

11.	Read back voltages	31

12.	Conclusion	33

# PERFORMANCE SUMMARY

Module Number #		VM301-35-SCA

Operating temperature	72K

Pixel rate 		=	294 kHz

Outputs		= 	16

Frame readout time	=	1.001s

Detector bias:

VRstUc	=	0.0V

VnUc		=	0.5V 		(0.7V)

VDetCom	=	0.5V		(0.7V)

Mean DC level	= 	3.285V @ 0.5V

3.299V @ 0.7V

Detector transimpedance 	= 4.387µV/e- (20 – 70% of full well)

(0.5V)				= 4.634µV/e- (up to ~ 20% of full well)

Dark generation	= 	0.80e-/s/pixel (mean) @0.5V detector bias

Operability 		= 	97.67% (>0, < 8e-/s)

Read noise

Temporal	=	18.7e- (rms)

Pixel-to-pixel	=	21.4e- (rms)

Full well (ke-) 		=	87ke- (@ 0.5V)

120ke- (@ 0.7V)

Non-linearity		=	3.5% @0.5V, 5.2% @ 0.7V

| QE* | Mean (%) | Sigma | Non-Uniformity (sigma / mean) |
| --- | --- | --- | --- |
| J-Band | 92.4 | 4.5 | 4.9% |
| H-Band | 96.3 | 4.6 | 4.8% |
| K-Band | 90.7 | 3.9 | 4.3% |

* - Includes only operable pixels

Persistence signal		=	7.8e-

Total Pixel operability 	=	97.3%

# DC Uniformity

DC Uniformity (0.7V detector bias)

Mean DC @0.5V = 3.285V, Sigma = 0.039V

Mean DC @0.7V = 3.299V, Sigma = 0.064V

DC Uniformity (Histogram)

Mean DC = 3.299V, Sigma = 0.064V

# System gain and conversion gain

System gain (20 -70% of full well) – 0.7V bias

System gain @ 0.5V bias = 5.197e-/ADU

System gain @ 0.7V bias = 4.864e-/ADU

System gain (up to ~20% full well) – 0.7V bias

System gain @ 0.5V bias = 4.921e-/ADU

System gain @ 0.7V bias = 4.559e-/ADU

Transimpedance conversion gain vs. signal level (0.7V)

Transimpedance conversion gain @ 0.5V bias 	= 4.387µV/e- (20 – 70% full well)

= 4.634µV/e- (up to ~20% full well)

Transimpedance conversion gain @ 0.7V bias	= 4.688µV/e- (20 – 70% full well)

= 5.0µV/e- (up to ~20% full well)

Linearity and full well

Photon transfer – full well obtained from noise roll-over in the plot (0.7V)

Full well @ 0.5V bias = 87ke-

Full well @ 0.7V bias = 120ke-

Non-linearity plot (0.7V)

Estimated from 10 – 80% of saturation levels

Non-linearity 	= ~ 3.5% (0.5V bias)

= ~ 5.2% (0.7V bias)

# Read noise

Read noise bitmap (temporal)

Estimated from 10 CDS frames (0.7V)

The array mean read noise (temporal) is 17.3e- (rms).

Number of pixels >32e- noise is 1.28%.

Read noise histogram (temporal)

Obtained from 10 CDS frames with minimum detector integration time

Histogram shows the array mean excluding the pixels >32e-

The array mean for all pixels in the frame is 18.7e-.

1.28% pixels show read noise >32e.

Pixel-to-pixel read noise (excluding pixels with dark > 8e/s)

Obtained from two subsequent dark frames with minimum detector integration time

Closely agrees with the temporal noise

Read noise measured the 1st output = 18.9e-.

# Dark generation

Dark generation (@ 0.7V detector bias)

Measured after a settling time of ~3hrs (i.e. after adjusting any detector bias voltage)

An average of 10 CDS frames (each with 240 sec. integration time) is used to estimate the mean dark generation. Mean dark generation for the full frame excluding pixels with 0 < dark generation > 8e-/sec is estimated as 0.81 e-/sec/pixel (σ = 0.46). The pixels in limits are ~97.5%.

Mean dark @0.5V bias = 0.80e-/s/pix with pixels in limits =97.7 % (with VnUc = 0.7V)

Mean dark @0.7V bias = 0.81e-/s/pix with pixels in limits =97.5 % (with VnUc = 0.7V)

Dark generation bitmap (0.7V detector bias) – mean dark = 0.81e-/s/pix. (0 < dark <8)

Dark from a region relatively free from hot pixels (in the centre) measured as 0.61e-/s/pix.

# Quantum Efficiency

QE bitmap (J-band)

Image stretched between 0.38 and 1.20 limits

Mean = 0.924

Sigma = 0.045

QE histogram (J-band)

Mean and sigma are obtained from operable pixels

Non-uniformity 	= 4.9% (for pixels in limits)

Operability = 98.34%

Blackbody temperature:	low	= 350ºC 	(9.83E8 ph/cm2/s)

High	= 380ºC 	(2.26E9 ph/cm2/s)

QE bitmap (H-band)

Mean = 0.963

Sigma = 0.046

Image stretched between 0.47 and 1.2 limits

Mean and sigma are obtained from operable pixels

QE histogram (H-band)

Mean and sigma are obtained from operable pixels

Non-uniformity 	= 4.8% (for pixels in limits)

Operability = 98.31%

Blackbody temperature:	low	= 220ºC 	(1.25E9 ph/cm2/s)

High	= 250ºC 	(3.38E9 ph/cm2/s)

QE bitmap in K-band

Mean = 0.907

Sigma = 0.039

Image stretched between 0.47 and 1.20 limits

Mean and sigma are obtained from operable pixels

QE histogram (K-band)

Mean and sigma are obtained from operable pixels

Non-uniformity 	= 4.3% (for pixels in limits)

Operability = 98.32%

Blackbody temperature:	low	= 120ºC 	(1.36E9 ph/cm2/s)

High	= 150ºC 	(4.31E9 ph/cm2/s)

# Persistence

Figure shows the histogram of the persistence signal in a 10s dark after a minute after saturation.

Persistence signal in a 10s dark

Figure shows the persistence signal in a 10s dark after 1 minute after saturation with ~200ke- fluence. The detector is continuously read every 10s after saturation and the persistence decay is obtained.

Persistence signal decay

The array mean is 7.8e-.

# Inoperable Pixels

Inoperable pixels bitmap (white pixels)

Total inoperable pixels = ~2.70%

Obtained from dark operability (0 < dark <8e/s), and QE operability in J, H and K bands

About 1.6% of these are in the large area defects

Open pixels = 1230

# Temporal Stability

Temporal stability (K-band)

Ratio of two identical flats (50% full well) obtained 5 days apart

Flats are divided by one another and a histogram of the resulted frame is shown above. Inoperable pixels (see bitmap above) have been excluded from the measurement.

Temporal stability in J band = 0.73% (~ 50% full well, one day apart)

Temporal stability in H band = 0.59% (~ 50% full well, one day apart)

Ratio of J to K flat

# Cosmetics

Large area defects in a 480 x 420 section

Large area defect in a 330 x 260 section

Chip-off on top and two small cluster defects in a 360 x 150 section

AR coating is fine

# Read back voltages

Warm (before thermal cycle)

| DC | Name | Value1(V) | Value2(V) | Current(mA) |
| --- | --- | --- | --- | --- |
| DC1 | vhiReset | 3.999 | 3.979 | 0.20 |
| DC2 | vloReset | 0.005 | 0.005 | - |
| DC3 | vIdle | -2.363 | -2.344 | 0.19 |
| DC4 | vSlew | -3.999 | -3.970 | 0.29 |
| DC5 | vhiRowEn | 5.000 | 4.966 | 0.34 |
| DC6 | vloRowEn | 0.996 | 1.025 | 0.29 |
| DC7 | vCas | 2.998 | 2.993 | - |
| DC8 | vpD | 3.999 | 3.970 | 0.29 |
| DC9 | vrstUc | 0.005 | 0.005 | - |
| DC10 | vpUc | 3.496 | 3.418 | 0.78 |
| DC11 | vnUc | 0.601 | 0.654 | 0.53 |
| DC12 | vnOut | 2.197 | 2.388 | 1.91 |
| DC13 | vdetCom | 0.698 | 0.698 | - |
| DC14 | vpOut | 9.741 | 9.546 | 1.95 |
| DC15 | vBiasRef | 3.452 | 3.447 | - |

Read back voltages – Cold

| DC | Name | Value1(V) | Value2(V) | Current(mA) |
| --- | --- | --- | --- | --- |
| DC1 | vhiReset | 3.999 | 3.989 | 0.10 |
| DC2 | vloReset | 0.005 | 0.005 | - |
| DC3 | vIdle | -2.363 | -2.354 | - |
| DC4 | vSlew | -3.999 | -3.984 | 0.15 |
| DC5 | vhiRowEn | 5.005 | 5.000 | - |
| DC6 | vloRowEn | 1.001 | 0.996 | - |
| DC7 | vCas | 2.998 | 2.993 | - |
| DC8 | vpD | 3.999 | 3.975 | 0.24 |
| DC9 | vrstUc | 0.005 | 0.005 | - |
| DC10 | vpUc | 3.496 | 3.428 | 0.68 |
| DC11 | vnUc | 0.601 | 0.654 | 0.53 |
| DC12 | vnOut | 2.197 | 2.378 | 1.81 |
| DC13 | vdetCom | 0.698 | 0.698 | - |
| DC14 | vpOut | 9.746 | 9.556 | 1.90 |
| DC15 | vBiasRef | 3.452 | 3.452 | - |

Read back voltages – warm (after thermal cycle)

| DC | Name | Value1(V) | Value2(V) | Current(mA) |
| --- | --- | --- | --- | --- |
| DC1 | vhiReset | 3.994 | 3.975 | 0.19 |
| DC2 | vloReset | 0.005 | 0.005 | - |
| DC3 | vIdle | -2.363 | -2.344 | 0.19 |
| DC4 | vSlew | -3.999 | -3.970 | 0.29 |
| DC5 | vhiRowEn | 5.000 | 4.966 | 0.34 |
| DC6 | vloRowEn | 0.996 | 1.025 | 0.29 |
| DC7 | vCas | 2.998 | 2.993 | - |
| DC8 | vpD | 3.999 | 3.965 | 0.34 |
| DC9 | vrstUc | 0.005 | 0.005 | - |
| DC10 | vpUc | 3.491 | 3.418 | 0.73 |
| DC11 | vnUc | 0.601 | 0.654 | 0.53 |
| DC12 | vnOut | 2.197 | 2.388 | 1.91 |
| DC13 | vdetCom | 0.698 | 0.698 | - |
| DC14 | vpOut | 9.741 | 9.541 | 2.00 |
| DC15 | vBiasRef | 3.452 | 3.447 | - |

# Conclusion

Module#35 shows very good performance except for the three large area defects. The detector has overall operability of 97.3%. Transimpedance conversion gain of 4.39µV/e- at 0.5V is measured which is slightly different from the RVS value (4.78µV/e-). The conversion gain is estimated both at 0.5V and 0.7V bias and also estimated for low and high signal levels. The low level conversion gains are applied for dark, read noise and persistence. The high signal conversion gain is applied for estimating the full well capacity. Full well of ~87ke- is obtained with 0.5V bias with non-linearity of 3.5%. Full well of ~120ke- is obtained at 0.7V detector bias with non-linearity ~5.2%. The high transimpedance conversion gain corresponds to a less diode junction capacitance and thus low well capacity.

A mean dark generation of 0.81e-/s/pix is measured for the array at 0.7V (0.80e-/s/pix @ 0.5V) after about of 3 hours settling time. Dark pixel operability is ~97.5% with pixels in limits (0 < dark < 8e-/s/pix).

High QE’s over 90% have been measured in all bands. RVS reported QEs in J (61%), H (80.0%) and K (79%) largely differ from ATC measurements. A note in the RVS data pack says that they had assumed the blackbody emissivity as 98% (different from previous modules?). That could explain some of these large differences.

Temporal read noise of ~17.3e- (rms) is obtained for the entire array while the pixel-to-pixel noise is ~19.8e-. RVS have reported the pixel-to-pixel noise in one output (channel-1) as 22.1e-. ATC measured the pixel-to-pixel noise in the same output as 20.5e-.

Persistence signal decays below 32e- in about 25s. The array mean persistence signal after a minute after saturation is 7.8e-.

This detector shows three large area defects (pixels are non-responsive) which are about 1.6% of total pixels. There is a fence of hot and high responsive pixels around each of these defects spreading few (~6) pixels.

Except for the large area defects, this detector shows very good performance and meets the specifications. High QEs with good uniformity in all bands, low dark generation, low persistence, and low noise make this detector a good science detector.

---oOo---